IEEE P802.15

Wireless Personal Area Networks

| Project | IEEE P802.15 Task Group for Wireless Personal Area Networks (WPANs) |
| --- | --- |
| Title | 802.15.6 Call for Applications - Response Summary |
| Date Submitted | 1 July 2008 |
| Source | Daniel Lewis (Editor) NICTA Canberra Research Laboratory 7 London Circuit Canberra ACT 2602 Australia | Voice: +61 2 6267 6252 Fax: +61 2 6267 6220 E-mail: daniel.lewis@nicta.com.au |
| Re: | Responses to Call for Applications in Body Area Networks |
| Abstract | This document seeks to summarize significant applications (those with more then a single slide of detail) presented to the IEEE 802.16.6 TG / IEEE 802.15-BAN SG / IEEE 802.15 IG in the context of the categorization scheme developed by the group and represented in the applications matrix. |
| Purpose | (1) To collate identified 802.15.6 applications into a single design document which will be distributed with the call for proposals. (2) To identify omitted or erroneous application information requiring updating before distribution with a call for proposals. (3) To provide a single point of reference (and application reference numbers) which can be used as a common point of reference when evaluating and comparing call for proposal responses. (4) Refine the applications matrix categorization scheme through use. (We have attempted to locate each significant presented application in the categorization scheme. Applications which appear out of place likely indicate a shortcoming in the categorization scheme which should be addressed through iteration and refinement of the document) |
| Notice | This document has been prepared to assist the IEEE P802.15. It is offered as a basis for discussion and is not binding on the contributing individual(s) or organization(s). The material in this document is subject to change in form and content after further study. The contributor(s) reserve(s) the right to add, amend or withdraw material contained herein. |
| Release | The contributor acknowledges and accepts that this contribution becomes the property of IEEE and may be made publicly available by P802.15. |

802.15.6 Proposed applications

Class A	 Medical applications	4

A1-1	 Wearable BAN (WMTS)	5

Electroencephalogram EEG	6

Electrocardiogram ECG	7

Electromyography EMG	8

Vital signals monitoring	9

Temperature (wearable)	12

Respiration monitor (wearable)	13

Heart rate monitor (wearable)	14

Pulse oximeter SpO2 wearable	15

Blood pressure monitor (wearable)	16

pH monitor (wearable)	17

Glucose sensor (wearable)	18

Hearing aid (ear to ear communication)	19

A1-1a	Disability assistance	21

Muscle tension monitor	22

Muscle tension stimulation	23

Weighing scale (wearable)	24

Fall detection (wearable)	25

A1-1b	Human performance management	26

Aiding professional and amature sport training	27

Assessing emergency service personnel performance	29

Assessing soldier fatigue and battle readiness	30

Non Human (Animal)	31

A1-2a	Implant BAN (MICS)	33

Glucose sensor (implant)	34

Cardiac arrhythmia monitor/recorder (implant)	36

Brain liquid pressure sensor (implant)	37

Endoscope capsule (gastrointestinal)	38

Drug delivery capsule	40

Deep brain stimulator (eg, Epilepsy, Parkinson's therrepy)	41

Cortical stimulator	42

Visual neuro-stimulator	43

Audio neuro stimulator	44

Brain-computer interface	45

A1-2b	Remote control of medical devices	46

Pacemaker	47

Implantable cardioverter defibrallitor ICD	48

Implanted actuator	50

Insulin pump	51

Hearing aid (wearable and implanted)	52

Retina implants	53

A2-1	 In hospital (in general)	54

A2-2	 Outside Hospital	55

Class B	 Non-Medical applications (including entertainment)	56

B1	 Real-time video streaming	57

Video streaming (among portable devices - Lower resolution)	58

Video streaming (portable device external displays - Higher resolution)	60

3D video	62

B2	 Real-time audio streaming	63

Voice communication for headsets (moNo, and Lower quality)	64

Music for headsets (stereo, and Higher quality)	66

Sound track/Music - 5.1 channel (Higher bandwidth)	68

B3	 Data file transfer	70

Data file transfer (Word, Excel, software programs and etc)	71

Image file transfer (digital camera, scanner and etc.)	73

Video file transfer (camcorder, multimedia player and etc.)	74

B4	 Stream transfer	75

Remote control of entertainment devices	76

Body motion capture/gesture recognition, position	78

Control signal from PC peripheral devices (e.g. mouse click)	80

Forgotten things monitor	81

Emotion	83

Smart Key	85

Identification	86

Vital sign and body information based entertainment service	87

B5	 Entertainment applications	88

Gaming applications using BAN	89

Social Networking using BAN	91

# Class A		Medical applications

# A1-1		Wearable BAN (WMTS)

| Electroencephalogram EEG | Z001 |
| --- | --- |
| SAR Safety | Highest | Radio range | <5m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Electrocardiogram ECG | Z002 |
| --- | --- |
| SAR Safety | Highest | Radio range | <5m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Electromyography EMG | Z003 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Small |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Vital signals monitoring | Z004 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Temperature (wearable) | Z005 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Respiration monitor (wearable) | Z006 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Heart rate monitor (wearable) | Z007 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Pulse oximeter SpO2 wearable | Z008 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Blood pressure monitor (wearable) | Z009 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| pH monitor (wearable) | Z010 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Glucose sensor (wearable) | Z011 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Hearing aid (ear to ear communication) | Z012 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

# A1-1a	Disability assistance

| Muscle tension monitor | Z013 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Muscle tension stimulation | Z014 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Weighing scale (wearable) | Z015 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Fall detection (wearable) | Z016 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

# A1-1b	Human performance management

| Aiding professional and amature sport training | Z014 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Assessing emergency service personnel performance | Z015 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <13 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Assessing soldier fatigue and battle readiness | Z016 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <14 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Non Human (Animal) | Z012 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

# A1-2a	Implant BAN (MICS)

| Glucose sensor (implant) | Z017 |
| --- | --- |
| SAR Safety | High | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Cardiac arrhythmia monitor/recorder (implant) | Z018 |
| --- | --- |
| SAR Safety | High | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Brain liquid pressure sensor (implant) | Z019 |
| --- | --- |
| SAR Safety | High | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Endoscope capsule (gastrointestinal) | Z020 |
| --- | --- |
| SAR Safety | High | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness | If possible | Market size / region / year | Large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body to out-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | SwalLowable |  |  |
| Data rate (per link) | > 2Mbps | Energy scavenging possible |  | Privacy | Medium |  |  |
| Number of devices per piconet) | <6 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Drug delivery capsule | Z021 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | 1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Deep brain stimulator (eg, Epilepsy, Parkinson's therrepy) | Z022 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Cortical stimulator | Z023 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Visual neuro-stimulator | Z024 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Audio neuro stimulator | Z025 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Brain-computer interface | Z026 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

# A1-2b	Remote control of medical devices

| Pacemaker | Z027 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In/On-body & On/In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Implantable cardioverter defibrallitor ICD | Z028 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In/On & On/In | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Implanted actuator | Z029 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Insulin pump | Z030 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In/On & On/In | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Hearing aid (wearable and implanted) | Z031 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On & In Body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Retina implants | Z032 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In Body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

# A2-1		In hospital (in general)

# A2-2		Outside Hospital

# Class B		Non-Medical applications (including entertainment)

# B1		Real-time video streaming

| Video streaming (among portable devices - Lower resolution) | Z033 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 10Mbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Video streaming (portable device external displays - Higher resolution) | Z034 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 10~20Mbps? | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| 3D video | Z035 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 100Mbps? | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

# B2		Real-time audio streaming

| Voice communication for headsets (moNo, and Lower quality) | Z36 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Music for headsets (stereo, and Higher quality) | Z37 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 1Mbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Sound track/Music - 5.1 channel (Higher bandwidth) | Z38 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 5Mbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

# B3		Data file transfer

| Data file transfer (Word, Excel, software programs and etc) | Z39 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Med | Form Factor |  |  |  |
| Data rate (per link) | 10Mbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Image file transfer (digital camera, scanner and etc.) | Z40 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 10Mbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Video file transfer (camcorder, multimedia player and etc.) | Z41 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 10Mbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

# B4		Stream transfer

| Remote control of entertainment devices | Z42 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery |  |  |  |
| Duty cycle (per device) % per time | Low | QoS - sensitive to latency | Yes | Power by: energy scavenging |  |  |  |
| Further information: |

| Body motion capture/gesture recognition, position | Z43 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | Med |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | Med | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Control signal from PC peripheral devices (e.g. mouse click) | Z44 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | Low | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Forgotten things monitor | Z45 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Symmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | Med | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Emotion | Z46 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | Low | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Smart Key | Z47 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Identification | Z48 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Vital sign and body information based entertainment service | Z49 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

# B5		Entertainment applications

| Gaming applications using BAN | Z50 |
| --- | --- |
| SAR Safety | Medium | Radio range | <5m | Setup time for a new link | <1s | Cost | Low |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness | Yes | Market size / region / year | Very large |
| Topology | Star, P2P | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-air to On-Body, On-body to In-air, On-body to On-body | Covered by Other Standards | Bluetooth |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | 200kbps ~2Mbps | Energy scavenging possible | Yes If possible | Privacy | Medium |  |  |
| Number of devices per piconet) | 12 to 24 | QoS - sensitive to error | Medium | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <30% | QoS - sensitive to latency | High | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Social Networking using BAN | Z51 |
| --- | --- |
| SAR Safety | Medium | Radio range | <1.5m | Setup time for a new link | <1s | Cost | Low |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness | No | Market size / region / year | Very large |
| Topology | P2P | Robustness /reliability high/med/low | Low | Channel (e.g. in/out body) | In-air to In-air | Covered by Other Standards | Bluetooth |
| Data link speed sym/ asymmetric | Symmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | <200Kbps | Energy scavenging possible | No | Privacy | High |  |  |
| Number of devices per piconet) | 2 | QoS - sensitive to error | Low | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency | Low | Power by: energy scavenging | No |  |  |
| Further information: |

# Acknowledgements

This document was compiled from the following presentations presented at the IEEE 802.16.6 TG / IEEE 802.15-BAN SG / IEEE 802.15-BAN IG.

Any omissions from this list is unintentional