IEEE P802.15

Wireless Personal Area Networks

| Project | IEEE P802.15 Task Group for Wireless Personal Area Networks (WPANs) |
| --- | --- |
| Title | 802.15.6 Call for Applications - Response Summary |
| Date Submitted | Revision 10: 1 July 2008 Revision 21 8 Sep 2008 Revision 2 9 Sep 2008 Revision 3 9 Sep 2008 Revision 4 2 Oct 2008 |
| Source | Daniel Lewis (Editor) NICTA Canberra Research Laboratory 7 London Circuit Canberra ACT 2602 Australia | Voice: +61 2 6267 6252 Fax: +61 2 6267 6220 E-mail: daniel.lewis@nicta.com.au |
| Re: | Responses to Call for Applications in Body Area Networks |
| Abstract | This document presents applications which the IEEE 802.15.6 TG requests standards proposals support. The document summarises the significant applications (those with more then a single slide of detail) presented to the IEEE 802.16.6 TG / IEEE 802.15-BAN SG / IEEE 802.15 IG. The applications are presented in the context of the categorization scheme developed by the group and originally represented in the applications matrix. |
| Purpose | (1) To identify applications which the IEEE 802.15.6 TG request standard proposals support (2) To collate all significant 802.15.6 applications into a single design document. (3) To provide a single point of reference (and application reference numbers) which can be used as a common point of reference when evaluating and comparing call for proposal responses. |
| Notice | This document has been prepared to assist the IEEE P802.15. It is offered as a basis for discussion and is not binding on the contributing individual(s) or organization(s). The material in this document is subject to change in form and content after further study. The contributor(s) reserve(s) the right to add, amend or withdraw material contained herein. |
| Release | The contributor acknowledges and accepts that this contribution becomes the property of IEEE and may be made publicly available by P802.15. |

802.15.6 Proposed applications

Class A	 Medical applications	4

A1-1	 Wearable BAN (WMTS)	5

Electroencephalogram EEG	6

Electrocardiogram ECG	7

Electromyography EMG	9

Vital signals monitoring	10

Temperature (wearable)	13

Respiration monitor (wearable)	15

Heart rate monitor (wearable)	16

Pulse oximeter SpO2 wearable	17

Blood pressure monitor (wearable)	19

pH monitor (wearable)	21

Glucose sensor (wearable)	22

Hearing aid (ear to ear communication)	23

A1-1a	Disability assistance	25

Muscle tension monitor	26

Muscle tension stimulation	27

Weighing scale (wearable)	28

Fall detection (wearable)	30

A1-1b	Human performance management	31

Aiding professional and amature sport training	32

Assessing emergency service personnel performance	34

Assessing soldier fatigue and battle readiness	35

Non Human (Animal)	36

A1-2a	Implant BAN (MICS)	38

Glucose sensor (implant)	39

Cardiac arrhythmia monitor/recorder (implant)	41

Brain liquid pressure sensor (implant)	42

Endoscope capsule (gastrointestinal)	43

Drug delivery capsule	45

Deep brain stimulator (eg, Epilepsy, Parkinson's therrepy)	46

Cortical stimulator	47

Visual neuro-stimulator	48

Audio neuro stimulator	49

Brain-computer interface	50

A1-2b	Remote control of medical devices	51

Pacemaker	52

Implantable cardioverter defibrallitor ICD	53

Implanted actuator	55

Insulin pump	56

Hearing aid (wearable and implanted)	57

Retina implants	58

A2-1	 In hospital (in general)	59

A2-2	 Outside Hospital	60

Class B	 Non-Medical applications (including entertainment)	61

B1	 Real-time video streaming	62

Video streaming (among portable devices - Lower resolution)	63

Video streaming (portable device external displays - Higher resolution)	65

3D video	67

B2	 Real-time audio streaming	68

Voice communication for headsets (moNo, and Lower quality)	69

Music for headsets (stereo, and Higher quality)	71

Sound track/Music - 5.1 channel (Higher bandwidth)	73

B3	 Data file transfer	75

Data file transfer (Word, Excel, software programs and etc)	76

Image file transfer (digital camera, scanner and etc.)	78

Video file transfer (camcorder, multimedia player and etc.)	79

B4	 Stream transfer	80

Remote control of entertainment devices	81

Body motion capture/gesture recognition, position	83

Control signal from PC peripheral devices (e.g. mouse click)	85

Forgotten things monitor	86

Emotion	88

Smart Key	90

Identification	91

Vital sign and body information based entertainment service	92

B5	 Entertainment applications	93

Gaming applications using BAN	94

Social Networking using BAN	96

B6	 Emergency (non-medical)	98

Emergency (non-medical)	99

Acknowledgements	100

# Class A		Medical applications

# A1-1		Wearable BAN

| Electroencephalogram EEG | Z001 |
| --- | --- |
| SAR Safety | Highest | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Electrocardiogram ECG | Z002 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <30s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness | No | Market size / region / year | Very large |
| Topology | Star or point to point | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | 3kbps per ECG channel | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: For example: 12 channel ECG (raw streaming): 250Hz sampling rate 12 bits / sample 3kbps per channel 36kbps total 50kbps with small protocol overhead 12-channel ECG (compressed example. References required) 250Hz sampling rate 12 bits / sample 2kbps total Some applications do not require a 12 channel ECG. Fewer channels result in proportionally lower bandwidth requirements. Other Requirements: Low Latency is required for certain care scenarios. Also some scenarios require retransmission of lost data. In this case low latency is not as much of an issue, but reliability is. |

| Electromyography EMG | Z003 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Small |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Vital signals monitoring | Z004 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <30s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Temperature (wearable) | Z005 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <30s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness | No | Market size / region / year | Very large |
| Topology | Star or point to point | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | < 16bps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: Two example scenarios: Continuous For example Sports/performance monitoring multiple times per minute. Occasional For example recording temperature during illness several times per day. Typical payload need only be two bytes per sample. Maximum sampling rate: 1Hz  16bps Sensor power consumption can be of the order of microwatts |

| Respiration monitor (wearable) | Z006 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Heart rate monitor (wearable) | Z007 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <30s | Cost | Low |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness | No | Market size / region / year | Large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | 1bps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |
| Example: One 10-bit reading per minute (in beats/min). Including protocol overheads (preamble, ECC, maybe ACK), this would be around 1bps. |  |

| Pulse oximeter SpO2 wearable | Z008 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <30s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star/tree/point to point | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | 32 bps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: Payload of approximately 4 bytes per reading Pulse rate (1 byte) Perfusion (2 bytes) Blood Oxygenation (1 byte) Data sent up to once per second Typical use cases: - Monitoring for 20 minutes per day - During a workout - Acute illness Power less than 50mW typically |

| Blood pressure monitor (wearable) | Z009 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | High | Form Factor | Medium |  |  |
| Data rate (per link) | < 10 bps | Energy scavenging possible | No | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | No |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: Based on a typical cuff-type blood pressure monitor  One reading taken at most, two to three times per day  3 bytes per reading  Not time critical  Assuming a cuff based blood pressure sensor, these devices are high power, e.g. 1.5W during pumping to inflate cuff |

| pH monitor (wearable) | Z010 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Glucose sensor (wearable) | Z011 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Hearing aid (ear to ear communication) | Z012 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

# A1-1a	Disability assistance

| Muscle tension monitor | Z013 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Muscle tension stimulation | Z014 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Weighing scale (wearable) | Z015 |
| --- | --- |
| SAR Safety | Medium | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree/point-to-point | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | <10 bps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: Approximately 2 bytes per reading Low power sensor of the order of a few 10s of mW if not less |

| Fall detection (wearable) | Z016 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

# A1-1b	Human performance management

| Aiding professional and amature sport training | Z014 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree/point-to-point | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | Varies per sensor type | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information:  Time synchronization between multiple motion sensors may be required in order to correlate data.  Both raw and compressed data can be sent over the link depending on the specific application. For example: node sends step rate or raw acceleration data. |

| Assessing emergency service personnel performance | Z015 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <13 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Assessing soldier fatigue and battle readiness | Z016 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <14 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <10% | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Non Human (Animal) | Z012 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

# A1-2a	Implant BAN

| Glucose sensor (implant) | Z017 |
| --- | --- |
| SAR Safety | High | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Cardiac arrhythmia monitor/recorder (implant) | Z018 |
| --- | --- |
| SAR Safety | High | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Very large |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Brain liquid pressure sensor (implant) | Z019 |
| --- | --- |
| SAR Safety | High | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year | Large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Endoscope capsule (gastrointestinal) | Z020 |
| --- | --- |
| SAR Safety | High | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness | If possible | Market size / region / year | Large |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body to out-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | SwalLowable |  |  |
| Data rate (per link) | > 2Mbps | Energy scavenging possible |  | Privacy | Medium |  |  |
| Number of devices per piconet) | <6 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Drug delivery capsule | Z021 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | 1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Deep brain stimulator (eg, Epilepsy, Parkinson's therrepy) | Z022 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Cortical stimulator | Z023 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Visual neuro-stimulator | Z024 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Audio neuro stimulator | Z025 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Brain-computer interface | Z026 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor | Small |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <50% | QoS - sensitive to latency | Yes | Power by: energy scavenging | If possible |  |  |
| Further information: |

# A1-2b	Remote control of medical devices

| Pacemaker | Z027 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In/On-body & On/In-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Implantable cardioverter defibrallitor ICD | Z028 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In/On & On/In | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Implanted actuator | Z029 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Insulin pump | Z030 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In/On & On/In | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Hearing aid (wearable and implanted) | Z031 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On & In Body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Retina implants | Z032 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | In Body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Extremely Low | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency |  | Power by: energy scavenging | If possible |  |  |
| Further information: |

# A2-1		In hospital (in general)

# A2-2		Outside Hospital

# Class B		Non-Medical applications (including entertainment)

# B1		Real-time video streaming

| Video streaming (among portable devices - Lower resolution) | Z033 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 10Mbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Video streaming (portable device external displays - Higher resolution) | Z034 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 10~20Mbps? | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| 3D video | Z035 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 100Mbps? | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

# B2		Real-time audio streaming

| Voice communication for headsets (moNo, and Lower quality) | Z36 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Music for headsets (stereo, and Higher quality) | Z37 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 1Mbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Sound track/Music - 5.1 channel (Higher bandwidth) | Z38 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 5Mbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error |  | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

# B3		Data file transfer

| Data file transfer (Word, Excel, software programs and etc) | Z39 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Med | Form Factor |  |  |  |
| Data rate (per link) | 10Mbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Image file transfer (digital camera, scanner and etc.) | Z40 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 10Mbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Video file transfer (camcorder, multimedia player and etc.) | Z41 |
| --- | --- |
| SAR Safety |  | Radio range | <3m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Relatively high | Form Factor |  |  |  |
| Data rate (per link) | 10Mbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | High | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

# B4		Stream transfer

| Remote control of entertainment devices | Z42 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery |  |  |  |
| Duty cycle (per device) % per time | Low | QoS - sensitive to latency | Yes | Power by: energy scavenging |  |  |  |
| Further information: |

| Body motion capture/gesture recognition, position | Z43 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | Med |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | Med | QoS - sensitive to latency | Yes | Power by: energy scavenging | No |  |  |
| Further information: |

| Control signal from PC peripheral devices (e.g. mouse click) | Z44 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | Low | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Forgotten things monitor | Z45 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star/tree | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Symmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | >24 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | Med | QoS - sensitive to latency |  | Power by: energy scavenging | No |  |  |
| Further information: |

| Emotion | Z46 |
| --- | --- |
| SAR Safety |  | Radio range | <5m | Setup time for a new link | <3s | Cost |  |
| Regulatory - Radio |  | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor |  |  |  |
| Data rate (per link) | 256kbps | Energy scavenging possible | Yes | Privacy | Low |  |  |
| Number of devices per piconet) | <12 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | Low | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Smart Key | Z47 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Identification | Z48 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

| Vital sign and body information based entertainment service | Z49 |
| --- | --- |
| SAR Safety |  | Radio range |  | Setup time for a new link |  | Cost |  |
| Regulatory - Radio |  | Coexistence |  | Location awareness |  | Market size / region / year |  |
| Topology |  | Robustness /reliability high/med/low |  | Channel (e.g. in/out body) |  | Covered by Other Standards |  |
| Data link speed sym/ asymmetric |  | Power consumption |  | Form Factor |  |  |  |
| Data rate (per link) |  | Energy scavenging possible |  | Privacy |  |  |  |
| Number of devices per piconet) |  | QoS - sensitive to error |  | Power by: battery |  |  |  |
| Duty cycle (per device) % per time |  | QoS - sensitive to latency |  | Power by: energy scavenging |  |  |  |
| Further information: |

# B5		Entertainment applications

| Gaming applications using BAN | Z50 |
| --- | --- |
| SAR Safety | Medium | Radio range | <5m | Setup time for a new link | <1s | Cost | Low |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness | Yes | Market size / region / year | Very large |
| Topology | Star, P2P | Robustness /reliability high/med/low | Medium | Channel (e.g. in/out body) | In-air to On-Body, On-body to In-air, On-body to On-body | Covered by Other Standards | Bluetooth |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | 200kbps ~2Mbps | Energy scavenging possible | Yes If possible | Privacy | Medium |  |  |
| Number of devices per piconet) | 12 to 24 | QoS - sensitive to error | Medium | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <30% | QoS - sensitive to latency | High | Power by: energy scavenging | If possible |  |  |
| Further information: |

| Social Networking using BAN | Z51 |
| --- | --- |
| SAR Safety | Medium | Radio range | <1.5m | Setup time for a new link | <1s | Cost | Low |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness | No | Market size / region / year | Very large |
| Topology | P2P | Robustness /reliability high/med/low | Low | Channel (e.g. in/out body) | In-air to In-air | Covered by Other Standards | Bluetooth |
| Data link speed sym/ asymmetric | Symmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | <200Kbps | Energy scavenging possible | No | Privacy | High |  |  |
| Number of devices per piconet) | 2 | QoS - sensitive to error | Low | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency | Low | Power by: energy scavenging | No |  |  |
| Further information: |

# B6		Emergency (non-medical)

| Emergency (non-medical) | Z52 |
| --- | --- |
| SAR Safety | High | Radio range | <5m | Setup time for a new link | <1s | Cost |  |
| Regulatory - Radio | Compliant | Coexistence | Yes | Location awareness |  | Market size / region / year |  |
| Topology | Star, Point-to Point | Robustness /reliability high/med/low | High | Channel (e.g. in/out body) | On-body | Covered by Other Standards |  |
| Data link speed sym/ asymmetric | Asymmetric | Power consumption | Low | Form Factor | Small |  |  |
| Data rate (per link) | < 10 kbps | Energy scavenging possible | Yes | Privacy | High |  |  |
| Number of devices per piconet) | 2 | QoS - sensitive to error | Yes | Power by: battery | Yes |  |  |
| Duty cycle (per device) % per time | <1% | QoS - sensitive to latency | Yes | Power by: energy scavenging | Yes |  |  |
| Further information: Example (1), Notification of a Fire in the home Example (2), Notification of flammable Gas, or poisonous gas in the home Sensors off the body, (eg built into the house) may detect a non-medical emergency event and need to urgently communicate this event to body worn devices, in order to notify the wearer of the emergency situation. |

# Notes:

# “Energy scavaingsavaging possible” refers to whether the power level of the sensor is low enough to be powered by energy scavagingsavaging technologies.  “power by energy scavenging” indicates whether a power scavenging approach would be appropriate in this applications context .

# Acknowledgements

This document was compiled from the following presentations presented at the IEEE 802.16.6 TG / IEEE 802.15-BAN SG / IEEE 802.15-BAN IG.

Any omissions from this list is unintentional