IEEE P802.15

Wireless Personal Area Networks

| Project | IEEE P802.15 Working Group for Wireless Personal Area Networks (WPANs) |
| --- | --- |
| Title | MedWiN Physical Layer Proposal |
| Date Submitted | 4 May, 2009 |
| Source | David Davenport1, Neal Seidl2, Jeremy Moss3, Maulin Patel4, Anuj Batra5, Jin-Meng Ho5, Srinath Hosur5, June Chul Roh5, Tim Schmidl5, Okundu Omeni6, Alan Wong6 GE Global Research, davenport@research.ge.com, 518-387-5041, 1 Research Circle, Niskayuna, NY, USA GE Healthcare, neal.seidl@med.ge.com, 414-362-3413, 8200 West Tower Avenue, Milwaukee, WI, USA Philips, j.moss@philips.com, +44 1223 427530, 101 Cambridge Science Park, Milton Road, Cambridge UK Philips, maulin.patel@philips.com, 914-945-6156, 345 Scarborough Road, Briarcliff Manor, NY, USA Texas Instruments, {batra@ti.com, 214-480-4220}, {jinmengho@ti.com, 214-480-1994}, {hosur@ti.com, 214-480-4432}, {jroh@ti.com, 214-567-4145}, {schmidl@ti.com, 214-480-4460}, 12500 TI Blvd, Dallas, TX, USA Toumaz Technology, {okundu.omeni@toumaz.com, +44 1235 438950}, {alan.wong@toumaz.com, +44 1235 438961}, Building 3, 115 Milton Park, Abingdon, Oxfordshire, UK |
| Re: | In response to TG6 Call for Proposal: 15-08-0811-03-0006-tg6-call-proposals.doc |
| Abstract | MedWiN Physical Layer Proposal |
| Purpose | Discussion |
| Notice | This document has been prepared to assist the IEEE P802.15. It is offered as a basis for discussion and is not binding on the contributing individual(s) or organization(s). The material in this document is subject to change in form and content after further study. The contributor(s) reserve(s) the right to add, amend or withdraw material contained herein. |
| Release | The contributor acknowledges and accepts that this contribution becomes the property of IEEE and may be made publicly available by P802.15. |

MedWiN Physical Layer Proposal

Authors:

David Davenport1, Neal Seidl2, Jeremy Moss3, Maulin Patel4
Anuj Batra5, Jin-Meng Ho5, Srinath Hosur5, June Chul Roh5
Tim Schmidl5, Okundu Omeni6, Alan Wong6

(1) 	GE Global Research, davenport@research.ge.com, 518-387-5041, 1 Research Circle, Niskayuna, NY, USA

(2) 	GE Healthcare, neal.seidl@med.ge.com, 414-362-3413, 8200 West Tower Ave, Milwaukee, WI, USA

(3) 	Philips, j.moss@philips.com, +44 1223 427530, 101 Cambridge Science Park, Milton Road, Cambridge, UK

(4)	Philips, maulin.patel@philips.com, +1-914-945-6156, 345 Scarborough Rd, Briarcliff Manor, NY, USA

(5) 	Texas Instruments, {batra, jinmengho, hosur, jroh, schmidl}@ti.com, 12500 TI Blvd, Dallas, TX 75243

(6) 	Toumaz Technology, {okundu.omeni, alan.wong}@toumaz.com, Bldg 3, 115 Milton Park, Abingdon, Oxfordshire, UK

May 4, 2009

Table of Contents

1	Acronyms and abbreviations	5

2	Physical-Layer Convergence Protocol Layer (PLCP)	6

2.1	Physical-Layer Protocol Data Unit	6

2.2	Data Rate-dependent Parameters for PLCP Header	7

2.3	Data Rate-dependent Parameters for PSDU	7

2.4	PLCP Preamble	9

2.5	PLCP Header	11

2.5.1	PHY Header	12

2.5.2	BCH Encoder for PLCP Header	13

2.5.3	Header Check Sequence	14

2.6	PSDU	15

2.6.1	Data Scrambler	16

2.6.2	BCH Encoder for PSDU	16

2.6.3	Pad Bits	19

2.6.4	Spreading	19

2.6.5	Bit Interleaver	20

2.6.6	Constellation Mapping	20

3	General Requirements	23

3.1	Operating Frequencies	23

3.1.1	Operating Frequency Bands	23

3.1.2	Channel Numbering	23

3.1.3	Preamble Sequence Assignment	24

3.2	PHY Layer Timing	24

3.2.1	Inter-frame Spacing	24

3.2.2	Receive-to-Transmit Turnaround Time	24

3.2.3	Transmit-to-Receive Turnaround Time	24

3.2.4	Time between Successive Transmissions	25

3.2.5	Center Frequency Switch Time	25

4	Transmitter Specifications	26

4.1	Transmit PSD Mask	26

4.2	Transmit Center Frequency Tolerance	26

4.3	Symbol Clock Frequency Tolerance	26

4.4	Clock Synchronization	26

4.5	Transmit Power	26

4.6	Transmitter Constellation Error	27

5	Receiver Specifications	28

5.1	Receiver Sensitivity	28

5.2	Adjacent Channel Rejection	28

5.3	Receiver Energy Detection	29

5.3.1	ED Threshold	29

5.3.2	ED Measurement Time	29

5.4	Receiver Clear Channel Assessment	29

Annex: Comparison Criteria	31

# Acronyms and abbreviations

BCH Code	Bose, Ray-Chaudhuri, Hocquenghem Code

DBPSK	Differential Binary Phase-Shift Keying

DQPSK	Differential Quadrature Phase-Shift Keying

D8PSK	Differential 8-Phase-Shift Keying

EIRP	Equivalent Isotropically Radiated Power

EVM	Error Vector Magnitude

FCS	Frame Check Sequence

HCS	Header Check Sequence

MAC	Media Access Control

PHY	Physical Layer

PLCP	Physical Layer Convergence Protocol

PPDU	Physical Layer Protocol Data Unit

PSDU	Physical Service Data Unit

SIFS	Short Inter-frame Spacing

SRRC 	Square-Root Raised Cosine

# Physical-Layer Convergence Protocol Layer (PLCP)

This clause provides a method for transforming a physical-layer service data unit (PSDU) into a physical-layer protocol data unit (PPDU). During the transmission, the PSDU shall be pre-appended with a physical-layer preamble and a physical-layer header in order to create the PPDU. At the receiver, the physical-layer preamble and physical-layer header serve as aids in the demodulation, decoding and delivery of the PSDU.

## Physical-Layer Protocol Data Unit

Figure 1 shows the format for the PPDU, which is composed of three main components: the PLCP preamble, the PLCP header and the PSDU. The components are listed in the order of transmission. The PLCP preamble is the first component of the PPDU (see Section 2.4). The purpose of the preamble is to aid the receiver in the timing synchronization and carrier-offset recovery.

The PLCP header is the second main component of the PPDU (see Section 2.5). The purpose of this component is to convey the necessary information about the PHY and MAC parameters to aid in decoding of the PSDU at the receiver.  The PLCP header can be further decomposed into a RATE field, a LENGTH field, a BURST MODE field, a header check sequence (HCS), BCH parity bits and reserved bits. The BCH parity bits are added in order to improve the robustness of the PLCP header. The PLCP header shall be transmitted using the given header data rate in the operating frequency band.

The PSDU is the last component of the PPDU (see Section 2.6). This component is formed by concatenating the MAC header with the MAC frame body and frame check sequence (FCS). The PSDU is then scrambled and optionally encoded by a BCH code. The PSDU shall be transmitted using any of the available data rates available in the operating frequency band.

Figure 1 — Standard PPDU structure

When transmitting the packet, the PLCP preamble is sent first, followed by the PLCP header and finally the PSDU.  All multiple byte fields shall be transmitted with least significant byte first and each byte shall be transmitted with least significant bit (LSB) first.

A compliant device shall be able to support transmission and reception in one of the following frequency bands: 2400 – 2483.5 MHz, 2360 – 2400 MHz, 402 – 405 MHz, 902 – 928 MHz, 950 – 956 MHz and 863 – 870 MHz.

## Data Rate-dependent Parameters for PLCP Header

Table 1 — Modulation parameters for PLCP Header

| Frequency Band (MHz) | Constellation | M | Symbol Rate (ksps) | Pulse Shape | Code Rate (k/n) | Spreading Factor (S) | Information Data Rate (kbps) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2360 – 2483.5 | /2-DBPSK | 2 | 631.58 | SRRC | 16/31 | 4 | 81.5 |
| 402 – 405 | /2-DBPSK | 2 | 176.47 | SRRC | 16/31 | 1 | 91.1 |
| 902 – 928 | /2-DBPSK | 2 | 315.79 | SRRC | 16/31 | 2 | 81.5 |
| 950 – 956 | /2-DBPSK | 2 | 250.00 | SRRC | 16/31 | 1 | 129.0 |
| 863 – 870 | /2-DBPSK | 2 | 125.00 | SRRC | 16/31 | 1 | 64.5 |

## Data Rate-dependent Parameters for PSDU

A compliant device shall be able to support transmission and reception in one of the following frequency bands. Support for /2-DBPSK and /4-DQPSK is mandatory, while support for /8-D8PSK is optional.

Table 2 — Modulation parameters for PSDU in 2400 MHz frequency band

| Frequency Band (MHz) | Constellation | M | Symbol Rate (ksps) | Pulse Shape | Code Rate (k/n) | Spreading Factor (S) | Information Data Rate (kbps) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2360 – 2483.5 | /2-DBPSK | 2 | 631.58 | SRRC | 51/63 | 4 | 127.8 |
|  | /2-DBPSK | 2 | 631.58 | SRRC | 51/63 | 2 | 255.6 |
|  | /2-DBPSK | 2 | 631.58 | SRRC | 51/63 | 1 | 511.3 |
|  | /4-DQPSK | 4 | 631.58 | SRRC | 51/63 | 1 | 1022.6 |

Table 3 — Modulation parameters for PSDU in 400 MHz frequency band

| Frequency Band (MHz) | Constellation | M | Symbol Rate (ksps) | Pulse Shape | Code Rate (k/n) | Spreading Factor (S) | Information Data Rate (kbps) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 402 – 405 | /2-DBPSK | 2 | 176.47 | SRRC | 45/63 | 1 | 126.1 |
|  | /4-DQPSK | 4 | 176.47 | SRRC | 45/63 | 1 | 252.1 |
|  | /4-DQPSK | 4 | 176.47 | SRRC | 1/1 | 1 | 352.9 |
|  | /8-D8PSK | 8 | 176.47 | SRRC | 51/63 | 1 | 428.6 |

Table 4 — Modulation parameters for PSDU in 900 MHz frequency band

| Frequency Band (MHz) | Constellation | M | Symbol Rate (ksps) | Pulse Shape | Code Rate (k/n) | Spreading Factor (S) | Information Data Rate (kbps) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 902 – 928 | /2-DBPSK | 2 | 315.79 | SRRC | 51/63 | 2 | 127.8 |
|  | /2-DBPSK | 2 | 315.79 | SRRC | 51/63 | 1 | 255.6 |
|  | /4-DQPSK | 4 | 315.79 | SRRC | 51/63 | 1 | 511.3 |
|  | /8-D8PSK | 8 | 315.79 | SRRC | 51/63 | 1 | 766.9 |
| 950 – 956 | /2-DBPSK | 2 | 250.00 | SRRC | 39/63 | 1 | 154.8 |
|  | /2-DBPSK | 2 | 250.00 | SRRC | 1/1 | 1 | 250.0 |
|  | /4-DQPSK | 4 | 250.00 | SRRC | 1/1 | 1 | 500.0 |
|  | /8-D8PSK | 8 | 250.00 | SRRC | 51/63 | 1 | 607.1 |
| 863 – 870 | /2-DBPSK | 2 | 125.00 | SRRC | 51/63 | 1 | 101.2 |
|  | /4-DQPSK | 4 | 125.00 | SRRC | 45/63 | 1 | 178.6 |
|  | /4-DQPSK | 4 | 125.00 | SRRC | 1/1 | 1 | 250.0 |
|  | /8-D8PSK | 8 | 125.00 | SRRC | 51/63 | 1 | 303.6 |

## PLCP Preamble

A preamble shall be added prior to the PLCP header in order to aid the receiver in packet detection, timing synchronization and carrier-offset recovery. Two unique preambles are defined in order to mitigate false alarms due to networks operating on adjacent channels. The mapping between channel number and preamble is defined in Section 3.1.3. Each preamble is constructed by concatenating a length-63 m-sequence with a 010101010 sequence. The former sequence can be used to implement packet detection, coarse-timing synchronization and carrier-offset recovery, while the latter sequence can be used to implement fine-timing synchronization.

The two preamble sequences are defined in Table 5 and Table 6. The preambles will be transmitted at the symbol rate for the desired band of operation and will be encoded using /2-DBPSK.

Table 5 — Preamble Sequence #1

| Bit | Bit Value | Bit | Bit Value | Bit | Bit Value | Bit | Bit Value |
| --- | --- | --- | --- | --- | --- | --- | --- |
| b0 | 0 | B18 | 0 | b36 | 1 | b54 | 0 |
| b1 | 1 | B19 | 1 | b37 | 0 | b55 | 0 |
| b2 | 0 | b20 | 0 | b38 | 0 | b56 | 0 |
| b3 | 1 | b21 | 0 | b39 | 1 | b57 | 1 |
| b4 | 0 | b22 | 1 | b40 | 0 | b58 | 1 |
| b5 | 1 | b23 | 0 | b41 | 1 | b59 | 1 |
| b6 | 1 | b24 | 0 | b42 | 0 | b60 | 1 |
| b7 | 0 | b25 | 1 | b43 | 0 | b61 | 1 |
| b8 | 0 | b26 | 1 | b44 | 0 | b62 | 1 |
| b9 | 1 | b27 | 1 | b45 | 1 | b63 | 0 |
| b10 | 1 | b28 | 0 | b46 | 1 | b64 | 1 |
| b11 | 0 | b29 | 0 | b47 | 0 | b65 | 0 |
| b12 | 1 | b30 | 0 | b48 | 0 | b66 | 1 |
| b13 | 1 | b31 | 1 | b49 | 0 | b67 | 0 |
| b14 | 1 | b32 | 0 | b50 | 0 | b68 | 1 |
| b15 | 0 | b33 | 1 | b51 | 1 | b69 | 0 |
| b16 | 1 | b34 | 1 | b52 | 0 | b70 | 1 |
| b17 | 1 | b35 | 1 | b53 | 0 | b71 | 0 |

Table 6 — Preamble Sequence #2

| Bit | Bit Value | Bit | Bit Value | Bit | Bit Value | Bit | Bit Value |
| --- | --- | --- | --- | --- | --- | --- | --- |
| b0 | 0 | B18 | 0 | b36 | 0 | b54 | 0 |
| b1 | 1 | B19 | 1 | b37 | 0 | b55 | 1 |
| b2 | 1 | b20 | 0 | b38 | 1 | b56 | 0 |
| b3 | 0 | b21 | 1 | b39 | 1 | b57 | 1 |
| b4 | 1 | b22 | 0 | b40 | 0 | b58 | 1 |
| b5 | 0 | b23 | 1 | b41 | 1 | b59 | 1 |
| b6 | 0 | b24 | 0 | b42 | 1 | b60 | 1 |
| b7 | 0 | b25 | 0 | b43 | 1 | b61 | 1 |
| b8 | 1 | b26 | 1 | b44 | 0 | b62 | 1 |
| b9 | 0 | b27 | 0 | b45 | 0 | b63 | 0 |
| b10 | 0 | b28 | 0 | b46 | 1 | b64 | 1 |
| b11 | 0 | b29 | 1 | b47 | 1 | b65 | 0 |
| b12 | 0 | b30 | 1 | b48 | 0 | b66 | 1 |
| b13 | 1 | b31 | 1 | b49 | 0 | b67 | 0 |
| b14 | 0 | b32 | 1 | b50 | 0 | b68 | 1 |
| b15 | 1 | b33 | 0 | b51 | 1 | b69 | 0 |
| b16 | 1 | b34 | 0 | b52 | 1 | b70 | 1 |
| b17 | 0 | b35 | 0 | b53 | 1 | b71 | 0 |

## PLCP Header

A PLCP header shall be added after the PLCP preamble to convey information about the PHY and MAC parameters that is needed at receiver in order to decode the PSDU.  The BCH encoded PLCP header shall be formed as shown in Figure 2:

- Format the PHY header as shown in Figure 2 based on information  provided by the MAC

- Calculate the 2-bit HCS value over the PHY header

- Apply a BCH (31,16) code to the concatenation of the PHY header (14 bits) and HCS (2 bits)

The resulting encoded bits are modulated using the appropriate parameters specified in Table 1 for the desired frequency band of operation.

Figure 2 — Block diagram of PLCP header construction

### PHY Header

The PHY header contains information about the data rate of the MAC frame body, the length of the MAC frame body (which does not include the MAC header or the FCS) and information about the next packet – whether it is being sent in a burst mode.

The PHY header field shall be composed of 14 bits, numbered from 0 to 13 as illustrated in Figure 3. Bits 0-2 shall encode the RATE field, which conveys the information about the type of modulation, the symbol rate, frequency deviation or pulse shape, the coding rate, and the spreading factor used to transmit the PSDU. Bits 3-10 shall encode the LENGTH field, with the least-significant bit being transmitted first. Bit 13 shall encode whether or not the packet is being transmitted in the burst (streaming) mode. All other bits which are not defined in this Clause shall be understood to be reserved for future use and shall be set to zero.

Figure 3 — PHY Header bit assignment

#### Data Rate (RATE)

Depending on the data rate (RATE), bits R0–R2 shall be set according to the values in Table 7.

Table 7 — Rate-dependent parameters

| R0 – R2 | Data Rate (kbps) 2360-2483.5 MHz | Data Rate (kbps) 402-405 MHz | Data Rate (kbps) 902-928 MHz | Data Rate (kbps) 950-956 MHz | Data Rate (kbps) 863-870 MHz |
| --- | --- | --- | --- | --- | --- |
| 000 | 127.8 | 126.1 | 127.8 | 154.8 | 101.2 |
| 001 | 255.6 | 252.1 | 255.6 | 250.0 | 178.6 |
| 010 | 511.3 | 352.9 | 511.3 | 500.0 | 250.0 |
| 011 | 1022.6 | 428.6 | 766.9 | 607.1 | 303.6 |
| 100 | Reserved | Reserved | Reserved | Reserved | Reserved |
| 101 | Reserved | Reserved | Reserved | Reserved | Reserved |
| 110 | Reserved | Reserved | Reserved | Reserved | Reserved |
| 111 | Reserved | Reserved | Reserved | Reserved | Reserved |

#### PLCP Length Field (LENGTH)

The PLCP length field shall be an unsigned 8-bit integer that indicates the number of un-coded information bytes in the MAC frame body (which does not include the MAC header or the FCS).

#### Burst Mode Field (BM)

The MAC shall set the burst mode (BM) bit, as defined in Table 8, to indicate whether the next packet is part of a packet “burst”, i.e. burst mode transmission. In burst mode, the inter-frame spacing shall be equal to a pSIFS (see Section 3.2.4).

Table 8 — Burst Mode field

| Burst Mode (BM) bit | Next Packet Status |
| --- | --- |
| 0 | Next packet is not part of burst |
| 1 | Next packet is part of burst |

### BCH Encoder for PLCP Header

The PLCP header shall use a systematic BCH (31, 16, t = 3) code to improve the robustness of the PLCP header. The generator polynomial for this BCH code is:

.

The parity bits are determined by computing the remainder polynomial :

,

where  is the message polynomial:

,

and  and  are elements of . The message polynomial  is created as follows:  is the first bit of the PHY header,  is the last bit of the PHY header,  is the first bit of the HCS and  is the last bit of the HCS. The order of the parity bits is as follows:  is the first parity bit transmitted,  is the second parity bit transmitted, and  is the last parity bit transmitted.

A shift-register implementation of the BCH encoding process is shown in Figure 4, with the additions over . After  has been inserted into the shift register, the switch shall be moved from the message polynomial input connection to the shift register output connection (right-to-left).

Figure 4 — Shift register implementation of a systematic BCH (31, 16) encoder

### Header Check Sequence

The PHY header shall be protected with a 2-bit (CRC-2) header check sequence (HCS). The HCS shall be the ones complement of the remainder generated by the modulo-2 division of the PHY header by the polynomial: . The HCS bits shall be processed in the transmit order. A schematic of the processing order is shown in Figure 5. The registers shall be initialized to all ones.

Figure 5 — Block diagram of a CRC-2 implementation

## PSDU

The PSDU is the last major component of the PPDU and shall be constructed as shown in Figure 6.

- Form the non-scrambled PSDU by pre-pending the 7-byte MAC header to the MAC frame body and appending a 2-byte FCS to the result.

- The resulting PSDU is scrambled according to Section 2.6.1.

- If the code rate (k/n) < 1, the scrambled output is:

  - Divided into blocks of messages, where each message may contain shortened bits, according to Section 2.6.2.1;

  - The resulting messages are then encoded into codewords using the appropriate BCH encoder to achieve the desired code rate: a BCH (63, 39), a BCH (63, 45) or a BCH (63, 51) code, according to Sections 2.6.2.2 - 2.6.2.4;

  - Finally, the shortened bits are removed from each of the codewords.

- Pad bits are then added in order to ensure alignment on a symbol boundary according to Section 2.6.3.

- If the spreading factor is 2 or 4, the resulting un-coded or coded bits are spread using a repetition code, according to Section 2.6.4,  and then interleaved using a bit interleaver defined in Section 2.6.5.

- Finally, the resulting un-coded or coded and potentially spread and interleaved bit steam is then mapped onto the appropriate constellation (see Section 2.6.6), which is determined by the data rate and frequency band of operation (see Table 2 through Table 4).

Figure 6 — Block diagram of PSDU construction

### Data Scrambler

The polynomial  shall be used to scramble the PSDU, i.e., the concatenation of the MAC header, MAC frame body and FCS. The feed-through configuration of the scrambler and descrambler is self-synchronizing, which requires no prior knowledge of the transmitter initialization of the scrambler for receive processing. Figure 7 and Figure 8 show typical implementations of the data scrambler and descrambler, but other implementations are possible.

Figure 7 — Block diagram of data scrambler

Figure 8 — Block diagram of data de-scrambler

### BCH Encoder for PSDU

The following coding rates: 39/63, 45/63 and 51/53, shall be supported by the systematic BCH encoders. The information bits shall be encoded using BCH encoding process defined in Section 2.6.2.1. The definitions for the three systematic BCH encoders shall be defined in Sections 2.6.2.2 – 2.6.2.4.

#### BCH Encoding Process

The scrambled PSDU shall be encoded using the following procedure:

- Compute the number of bits in the PSDU :

where   is the number of bytes in the MAC header,  is the number of bytes in the MAC frame body and  is the number of bytes in the FCS.

- Calculate the number of BCH codeword :

where   is the number of message bits for the selected BCH code.

- Compute the number of shortening bits, , to be padded to the  data bits before encoding as follows:

- The shortening bits shall be equally distributed over all  codewords with the first  codewords being shortened one bit more than the remaining codewords. Let

.

Thus, the first  codewords will have  shortened bits (message bits that are set to 0), while the remaining codewords will have  shortened bits. After encoding, the shortened bits shall be discard prior to transmission, i.e., the shortened bits are never transmitted on-air.

The BCH encoding process is illustrated in Figure 9 for a single codeword.

Figure 9 — BCH encoding process for a single codeword

#### BCH (63, 39) Encoder

The generator polynomial for a systematic BCH (63, 39, t = 4) code is given by:

.

The parity bits are determined by computing the remainder polynomial :

,

where  is the message polynomial:

,

and  and  are elements of . The message polynomial  is created as follows:  is the first bit of the message and  is the last bit of the message, which maybe a shortened bit. The order of the parity bits is as follows:  is the first parity bit transmitted,  is the second parity bit transmitted, and  is the last parity bit transmitted.

#### BCH (63, 45) Encoder

The generator polynomial for a systematic BCH (63, 45, t = 3) code is given by:

.

The parity bits are determined by computing the remainder polynomial :

,

where  is the message polynomial:

,

and  and  are elements of . The message polynomial  is created as follows:  is the first bit of the message and  is the last bit of the message, which maybe a shortened bit. The order of the parity bits is as follows:  is the first parity bit transmitted,  is the second parity bit transmitted, and  is the last parity bit transmitted.

#### BCH (63, 51) Encoder

The generator polynomial for a systematic BCH (63, 51, t = 2) code is given by:

.

The parity bits are determined by computing the remainder polynomial :

,

where  is the message polynomial:

,

and  and  are elements of . The message polynomial  is created as follows:  is the first bit of the message and  is the last bit of the message, which maybe a shortened bit. The order of the parity bits is as follows:  is the first parity bit transmitted,  is the second parity bit transmitted, and  is the last parity bit transmitted.

### Pad Bits

Pad bits shall be appended after the BCH encoder to ensure that the bit stream aligns on a symbol boundary. The number of pad bits, , that shall be inserted is a function of the number of PSDU bits , the number of codewords , the number of parity bits , and the modulation constellation size :

.

The pad bits shall be appended to scrambled and encoded PSDU and all of the appended pad bits shall be set to 0.

### Spreading

For a spreading factor of 2, each input bit is repeated two times (see illustration in Figure 10a). For a spreading factor of 4, each input bit is repeated four times (see illustration in Figure 10b).

Figure 10 — Spreading scheme

### Bit Interleaver

The output of the spreader shall be interleaved prior to modulation to provide robustness against error propagation. The exact structure of the bit interleaver depends on the number of un-coded or coded bits that will be transmitted on-air, which is given by:

.

If , the bit interleaving operation is performed by first grouping the spread bits into blocks of 2S bits, where S is the spreading factor, and then using a block interleaver of size S(2 to permute the bits. Let the sequences  and , where , represent the input and output bits of the S(3 bit interleaver, respectively. The output of the S(2 bit interleaver is given by the following relationship:

.

If , the bit interleaving operation is performed by grouping the first 3S spread bits into a single block and then using a block interleaver of size S(3 to permute the bits within that single block. Let the sequences  and , where , represent the input and output bits of the S(3 bit interleaver, respectively. The output of the S(3 bit interleaver is given by the following relationship:

.

The remaining spread bits are then grouped into blocks of 2S bits and interleaved using the block interleaver of size S(2 shown in (19).

### Constellation Mapping

The un-coded or coded, potentially spreading and interleaved bit stream shall be mapped onto one of three rotated and differentially-encoded constellations: /2-DBPSK, /4-DQPSK, or /8-D8PSK. The encoded information is carried in the phase transitions between symbols. For the PLCP preamble to PLCP header transition, the phase change is relative to the last symbol for the PLCP preamble. For the PLCP header to PSDU transition, the phase change is relative to the last symbol for the PLCP header.

The binary bit stream  shall be mapped onto a corresponding complex-values sequence  as follows:

.

where and the relationship between the bit stream  and the phase change  is given in Table 9 through Table 11 for /2-DBPSK, /4-DQPSK, or /8-D8PSK, respectively.

Table 9 — /2-DBPSK Mapping

| b(n) | k |
| --- | --- |
| 0 | /2 |
| 1 | 3/2 |

Table 10 — /4-DQPSK Mapping

| b(2n) | b(2n+1) | k |
| --- | --- | --- |
| 0 | 0 | /4 |
| 0 | 1 | 3/4 |
| 1 | 0 | 7/4 |
| 1 | 1 | 5/4 |

Table 11 — /8-D8PSK Mapping

| b(3n) | b(3n+1) | b(3n+2) | k |
| --- | --- | --- | --- |
| 0 | 0 | 0 | /8 |
| 0 | 0 | 1 | 3/8 |
| 0 | 1 | 0 | 7/8 |
| 0 | 1 | 1 | 5/8 |
| 1 | 0 | 0 | 15/8 |
| 1 | 0 | 1 | 13/8 |
| 1 | 1 | 0 | 9/8 |
| 1 | 1 | 1 | 11/8 |

# General Requirements

## Operating Frequencies

### Operating Frequency Bands

A compliant device shall be able to support transmissions and reception in one of the following frequency bands:

- 2400 – 2483.5 MHz

- 2360 – 2400 MHz

- 402 – 405 MHz

- 902 – 928 MHz

- 950 – 956 MHz

- 863 – 870 MHz

### Channel Numbering

The relationship between center frequency, fc, and channel number, nc, is shown in Table 12.

Table 12 — Relationship between Center Frequency and Channel Number

| Frequency Band (MHz) | Relationship between fc and nc |
| --- | --- |
| 2400 – 2483.5 | fc = 2402.00 + 1.00 ( nc (MHz), nc = 0, …, 78 |
| 2360 – 2400 | fc = 2362.00 + 1.00 ( nc (MHz), nc = 0, …, 37 |
| 402 – 405 | fc = 402.15 + 0.30 ( nc (MHz), nc = 0, …, 9 |
| 902 – 928 | fc = 903.50 + 0.50 ( nc (MHz), nc = 0, …, 47 |
| 950 – 956 | fc = 951.10 + 0.40 ( nc (MHz), nc = 0, …, 11 |
| 863 – 870 | fc = 865.60 + 0.20 ( g(nc) (MHz), nc = 0, …, 14 |

The mapping function g(nc) used in the 863 – 870 MHz frequency band is defined as follows:

,

### Preamble Sequence Assignment

The relationship between channel number, nc, and preamble sequence used in the PLCP preamble is shown in Table 13.

Table 13 — Relationship between Channel Number and Preamble

| Preamble Sequence | Relationship between preamble sequence and nc |
| --- | --- |
| 1 | rem(nc,2) = 0 |
| 2 | rem(nc,2) = 1 |

## PHY Layer Timing

The values for the PHY layer timing parameters are defined in Table 14.

Table 14 — Physical layer timing parameters

| PHY Parameter | Value |
| --- | --- |
| pSIFS | 20 s |
| pCCATime | 63 preamble symbols |
| pChannelSwitchTime | 50 s |

### Inter-frame Spacing

The inter-frame spacing parameters are given in Table 15.

Table 15 — Inter-frame spacing parameters

| MAC Parameter | Value |
| --- | --- |
| SIFS | pSIFS |

### Receive-to-Transmit Turnaround Time

The RX-to-TX turnaround time shall not be greater than pSIFS. This turnaround time shall be measured at the air interface. The time elapsed from the last received symbol, to the first transmitted symbol of the PLCP preamble for the next frame shall not be greater than pSIFS.

### Transmit-to-Receive Turnaround Time

The TX-to-RX turnaround time shall not be greater than pSIFS. This turnaround time shall be measured at the air interface. The time elapsed from the last transmitted symbol until the receiver is ready to begin the reception of the next PHY frame shall not be greater than pSIFS.

### Time between Successive Transmissions

For burst mode transmissions, the inter-frame spacing between uninterrupted successive transmissions by a device shall be fixed to exactly pSIFS. The inter-frame spacing time shall be measured at the air interface. The time elapsed from the leading edge of the last transmitted symbol to the leading edge of the first transmitted symbol of the PLCP preamble for the following packet shall be fixed to exactly pSIFS.

### Center Frequency Switch Time

The center frequency switch time is defined as the interval from when the PHY transmits or receives the last valid symbol on one center frequency until it is ready to transmit or receive the next symbol on a different center frequency. The center frequency switch time shall not exceed pChannelSwitchTime.

# Transmitter Specifications

## Transmit PSD Mask

The transmitted spectral mask shall be less than -20 dBr (dB relative to the maximum spectral density of the signal) for  , where fc is channel center frequency and fBW is the channel bandwidth and is a function of the frequency band of operation and is defined in Table 16.

Table 16 — Channel bandwidth as a function of the frequency band of operation

| Frequency (MHz) | fBW |
| --- | --- |
| 2360 – 2483.5 | 1 MHz |
| 402 – 405 | 300 kHz |
| 902 – 928 | 500 kHz |
| 950 – 956 | 400 kHz |
| 863 – 870 | 200 kHz |

The transmitted spectral density also shall comply with all regulations defined by local regulatory bodies.

## Transmit Center Frequency Tolerance

The transmitted center frequency tolerance shall be (20 ppm maximum.

## Symbol Clock Frequency Tolerance

The symbol clock frequency tolerance shall be (20 ppm maximum.

## Clock Synchronization

The transmit center frequencies and the symbol clock frequency shall be derived from the same reference oscillator.

## Transmit Power

A transmitter shall be capable of transmitting at least –10 dBm EIRP in all frequency bands, except for 402 – 405 MHz, where a transmitter shall be capable of transmitting at most -16 dBm EIRP. Devices should transmit lower power when possible in order to reduce interference to other devices and systems.

The maximum transmit power is limited by local regulatory bodies.

## Transmitter Constellation Error

The modulation accuracy of the transmitter is determined via an error-vector magnitude (EVM) measurement, which is calculated over N baud-spaced received complex values . A decision is made for each received complex value. The ideal position of the chosen symbol is represented by the vector . The error vector  is defined as the distance from the ideal position to the actual position of the received complex values, i.e., .

The EVM is defined as shown in the equation below:

,

where S is the magnitude of the vector to the ideal constellation point. A transmitter shall have EVM values less than or equal to those listed in Table 17 when measure for N = TBD symbols. The EVM shall be measured on baseband I and Q samples after the received signal is passed through a reference receiver, which shall perform the following operations: carrier-frequency offset estimation and symbol timing recovery while making the measurements.

Table 17 — Permissible EVM numbers as a function of constellation size

| Constellation | EVM Error |
| --- | --- |
| /2-DBPSK | TBD % |
|  /4-DQPSK | TBD % |
|  /8-D8PSK | TBD % |

# Receiver Specifications

## Receiver Sensitivity

For a packet error rate (PER) of less than 10% with a PSDU of 255 bytes, the minimum receiver sensitivity numbers in AWGN for the highest data rate in each operating frequency band shall be the numbers listed in Table 18. The minimum input levels are measured at the antenna connector, where a noise figure of 10 dB (referenced at the antenna) and an implementation loss of 6 dB have been assumed.

Table 18 — Receiver Sensitivity Numbers

| Frequency Band (MHz) | Information Data Rate (kbps) | Minimum Sensitivity (dBm) |
| --- | --- | --- |
| 2360 – 2483.5 | 1022.6 | -86 |
| 402 – 405 | 425.6 | -87 |
| 902 – 928 | 766.9 | -84 |
| 950 – 956 | 607.1 | -85 |
| 863 – 870 | 303.6 | -88 |

## Adjacent Channel Rejection

The adjacent channel rejection shall be measured by setting the desired signal's strength 3 dB above the rate-dependent sensitivity specified in Table 18 and raising the power of the interfering signal until a 10% PER is reached  for a PSDU length of 255 bytes. The power difference between the interfering and the desired channel is the corresponding adjacent channel rejection. The interfering signal in the adjacent channel shall be a conformant PHY signal at the same information data rate, unsynchronized with the signal in the channel under test. For a compliant PHY, the corresponding adjacent channel rejection shall be no less than the numbers specified in Table 19.

Table 19 — Receiver Adjacent Channel Rejection Numbers

| Frequency Band (MHz) | Information Data Rate (kbps) | Adjacent Channel Rejection (dB) |
| --- | --- | --- |
| 2360 – 2483.5 | 1022.6 | 7 |
| 402 – 405 | 425.6 | 2 |
| 902 – 928 | 766.9 | 2 |
| 950 – 956 | 607.1 | 2 |
| 863 – 870 | 303.6 | 2 |

## Receiver Energy Detection

The receiver energy detection (ED) measurement is an estimate of the received signal power within the bandwidth of the channel. It is intended for use by a network/MAC layer as part of a channel selection algorithm. No attempt is made to identify or decode signals on the channel.

### ED Threshold

The minimum ED value (zero) shall indicate received power less than either

- 10 dB above the specified receiver sensitivity (see Section 5.1) OR

- that which is prescribed by local regulatory requirements,

which ever is lower.

The range of received power spanned by the ED values shall be at least 40 dB. Within this range, the mapping from the received power in decibels to ED value shall be linear with an accuracy of ± 6 dB.

### ED Measurement Time

The ED measurement time, to average over, shall be either

- 8 preamble symbol periods OR

- that which is prescribed by local regulatory requirements,

which ever is longer in duration.

## Receiver Clear Channel Assessment

The PHY shall provide the capability to perform CCA according to at least one of the following three methods:

CCA Mode 1: Energy above threshold. CCA shall report a busy medium upon detecting any energy above the ED threshold.

CCA Mode 2: Carrier sense only. CCA shall report a busy medium only upon the detection of a signal compliant with this standard with the same modulation and characteristics of the PHY that is currently in use by the device. This signal may be above or below the ED threshold. The CCA detection time shall be equal to pCCATime.

CCA Mode 3: Carrier sense with energy above threshold. CCA shall report a busy medium using a logical combination of

  - Detection of a signal with the modulation and characteristics of this standard and

  - Energy above the ED threshold, where the logical operator may be AND or OR.

The CCA parameters are subject to the following criteria:

a) The ED threshold shall correspond to a received signal power as prescribed in section 5.3.1.

b) The CCA detection time shall be equal to pCCATime (see Section 3.2).

Annex: Comparison Criteria

- Regulatory – Compliance to TG6 regulatory document (08-0034)

The proposed PHY complies with all local regulations, but not limited to:

- FCC 47 CFR Part 15

- FCC 47 CFR Part 95

- ETSI EN 300 440-1 v1.5.1

- ETSI EN 301 839-1 v1.2.1

- ETSI EN 300 220-1 v2.1.1

- ARIB STD-T96 Ver 1.0

- Raw PHY data rate (node to node)

The proposed PHY support scalable information data rates from 100 kbps – 1000 kbps.

- Transmission distance between two nodes

The proposed PHY supports a minimum distance of 3 meters for all data rates. See link budget section for more details.

- Packet error rate (PER)

The PER performance for an AWGN channel is shown in Figure 11 through Figure 13 as a function of SNR and the constellation size and coding rate, which corresponds to data rate.  All simulations were performed over 20000 packets with a PSDU = 256 bytes. The assumptions for the simulations included ideal packet detection, ideal header decoding, ideal symbol and zero carrier-frequency offset.

Figure 11 — PER curves for /2-DBPSK and various code rates

Figure 12 — PER curves for /4-DQPSK and various code rates

Figure 13 — PER curves for /8-D8PSK and various code rates

- Link Budget and Reliability

Assumptions: AWGN, 0 dBi TX and RX antennas, free-space path loss. Link budget table for 2360 – 2483.5 MHz:

Assumptions: AWGN, 0 dBi TX and RX antennas, free-space path loss, path loss inside the body value stated in A. J. Johansson, “Wireless communication with medical implants: Antenna and propagation,”  ISSN 1402-8662, 2004. Link budget table for 402 – 405 MHz:

Assumptions: AWGN, 0 dBi TX and RX antennas, free-space path loss. Link budget table for 902 – 928 MHz:

Assumptions: AWGN, 0 dBi TX and RX antennas, free-space path loss. Link budget table for 950 – 956 MHz:

Assumptions: AWGN, 0 dBi TX and RX antennas, free-space path loss. Link budget table for 863 – 870 MHz:

The channel fading statistics were calculated using the CM4 (on-body to external device) channel models. The channel gains were averaged over all orientations (0º, 90º, 180º, 270º). The transmitter was located on the chest and the patient was standing. The assumed velocity was 1 km/hr. In the following table, the free-space path loss was removed from the gain, since it is already incorporated into the AWGN link budget tables.

The results for the 90%-tile, 95%-tile and 99%-tile fade depth at 3 meters are given in Table 20.

Table 20 — Channel Fading Statistics

| Frequency Band (MHz) | 90% Fade Depth at 3 meters | 90% Fade Depth at 3 meters | 90% Fade Depth at 3 meters |
| --- | --- | --- | --- |
| 2360 – 2483.5 | 17.1 dB | 17.5 dB | 19.0 dB |
| 902 – 928 | 18.8 dB | 19.0 dB | 19.5 dB |
| 950 – 956 | 18.6 dB | 18.7 dB | 19.2 dB |
| 863 – 870 | 19.4 dB | 19.5 dB | 20.0 dB |

Table 21, which is the link budget in realistic channel environments, shows that the proposed PHY has sufficient link margin at 3 meters, even with 99% fade depth, and therefore, can provide an operating range of at least 3 meters for all data rates.

Table 21 — Link Budget in Realistic Channel Environments

| Parameter | Value | Value | Value | Value | Value | Value | Value | Value |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Center Frequency (fc) [MHz] | 2480 | 2480 | 928 | 928 | 956 | 956 | 870 | 870 |
| Data Rate (Rb) [kbps] | 127.8 | 1022.6 | 127.8 | 766.9 | 154.8 | 607.1 | 101.2 | 303.6 |
| AWGN Link Margin [dB] | 35.2 | 25.9 | 44.8 | 31.4 | 44.2 | 32.4 | 46.9 | 36.2 |
| 99% Fade Depth at 3 meters | 19.0 | 19.0 | 19.5 | 19.5 | 19.2 | 19.2 | 20.0 | 20.0 |
| Link Margin [dB] | 16.2 | 6.9 | 25.3 | 11.9 | 25.0 | 13.2 | 26.9 | 16.2 |

- Power Emissions Level

The proposed PHY meets at local regulations, including regulations on SAR.

- Interference and coexistence

This document specifies the receiver adjacent channel rejection numbers in Table 19.  These number specify the power levels, with respect to the power of the desired device, that can be tolerated on an adjacent channel and still satisfy a PER of 10% with a PSDU of 256 bytes, given a 3 dB margin above sensitivity.

Table 22 lists the required distance and frequency separation from the interfering device that is required in order to maintain a PER of 10% with a PSDU of 256 bytes, given a 6 dB margin above sensitivity.

Table 22 — Distance and Frequency Separation Required from Interfering Devices

| Interferer | Value |
| --- | --- |
| IEEE 802.11g @ 2.4 GHz, Ptx = +15 dBm | dint ≤ 8.0 m (↓), fsep ≥ 22 MHz (↑) |
| Bluetooth @ 2.4 GHz, Ptx = 0 dbm | dint ≤ 0.3 m (↓), fsep ≥ 2 MHz (↑) |

- Multiple Co-located Network Support

Each of the proposed frequency bands supports a minimum of 10 channels:

- 2400 – 2483.5 MHz: 79 channels

- 2360 – 2400 MHz: 38 channels

- 402 – 405 MHz: 10 channels

- 902 – 928 MHz: 48 channels

- 950 – 956 MHz: 12 channels

- 863 – 870 MHz: 15 channels

Multiple co-located networks can be supported via frequency-division multiple access (FDMA).

- Scalability

Four data rates are provided within each frequency band.

- Power Efficiency

| Frequency Band (MHz) | Data Rate (kbps) | TX | RX | Standby | Deep Sleep |
| --- | --- | --- | --- | --- | --- |
| 2360 – 2483.5 | 1022.6 | 2.9 mW | 3.1 mW | 50 W | 250 nW |
| 902 – 928 950 – 956 863 – 870 | 733.1 607.1 303.6 | 2.2 mW | 2.5 mW | 50 W | 250 nW |
| 402 – 405 | 428.6 | 1.9 mW | 2.1 mW | 50 W | 125 nW |