March, 2012	 IEEE P802.15-11-0836-06-0004j

IEEE P802.15

Wireless Personal Area Networks

| Project | IEEE P802.15 Working Group for Wireless Personal Area Networks (WPANs) |
| --- | --- |
| Title | TG4j Amendment Draft |
| Date Submitted | [15th March 2012] |
| Source | [Dave Evans] [Philips, 101 Cambridge Science Park, Cambridge, UK] | Voice: [+44 1293 886490] Fax: [ ] E-mail: [david.evans@philips.com] |
| Re: | [] |
| Abstract | [This is the draft of the TG4j Amendment] |
| Purpose | [To enable editing of the draft TG4j MBAN amendment.] |
| Notice | This document has been prepared to assist the IEEE P802.15. It is offered as a basis for discussion and is not binding on the contributing individual(s) or organization(s). The material in this document is subject to change in form and content after further study. The contributor(s) reserve(s) the right to add, amend or withdraw material contained herein. |
| Release | The contributor acknowledges and accepts that this contribution becomes the property of IEEE and may be made publicly available by P802.15. |

IEEE Standard for

Local and metropolitan area networks–

Part 15.4: Low-Rate Wireless Personal Area Networks (LR-WPANs)

Amendment: Alternative Physical Layer Extension to support Medical Body Area Networks (MBANS) services operating in the 2360-2400 MHz band

Sponsor

LAN/MAN Standards Committee
of the
IEEE Computer Society

IEEE-SA Standards Board

Prepared by the IEEE 802.15 Working Group for Wireless Personal Area Networks

Committee IEEE802

Copyright © 2012 by the Institute of Electrical and Electronics Engineers, Inc.

Three Park Avenue

New York, New York 10016-5997, USA

All rights reserved.

This document is an unapproved draft of a proposed IEEE Standard. As such, this document is subject to change. USE AT YOUR OWN RISK! Because this is an unapproved draft, this document must not be utilized for any conformance/compliance purposes. Permission is hereby granted for IEEE Standards Committee participants to reproduce this document for purposes of international standardization consideration. Prior to adoption of this document, in whole or in part, by another standards development organization, permission must first be obtained from the IEEE Standards Association Department (stds.ipr@ieee.org). Other entities seeking permission to reproduce this document, in whole or in part, must also obtain permission from the IEEE Standards Association Department.

IEEE Standards Association Department

445 Hoes Lane

Piscataway, NJ 08854, USA

Abstract: This amendment defines a physical layer for IEEE 802.15.4 in the 2360 to 2400 MHz band which complies with Federal Communications Commission (FCC) MBANS rules. This amendment defines modifications to the MAC needed to support this new physical layer.

Keywords: MBANS, medical body area network

IEEE Standards documents are developed within the IEEE Societies and the Standards Coordinating Committees of the IEEE Standards Association (IEEE-SA) Standards Board. The IEEE develops its standards through a consensus development process, approved by the American National Standards Institute, which brings together volunteers representing varied viewpoints and interests to achieve the final product. Volunteers are not necessarily members of the Institute and serve without compensation. While the IEEE administers the process and establishes rules to promote fairness in the consensus development process, the IEEE does not independently evaluate, test, or verify the accuracy of any of the information or the soundness of any judgments contained in its standards.

Use of an IEEE Standard is wholly voluntary. The IEEE disclaims liability for any personal injury, property or other damage, of any nature whatsoever, whether special, indirect, consequential, or compensatory, directly or indirectly resulting from the publication, use of, or reliance upon this, or any other IEEE Standard document.

The IEEE does not warrant or represent the accuracy or content of the material contained herein, and expressly disclaims any express or implied warranty, including any implied warranty of merchantability or fitness for a specific purpose, or that the use of the material contained herein is free from patent infringement. IEEE Standards documents are supplied “AS IS.”

The existence of an IEEE Standard does not imply that there are no other ways to produce, test, measure, purchase, market, or provide other goods and services related to the scope of the IEEE Standard. Furthermore, the viewpoint expressed at the time a standard is approved and issued is subject to change brought about through developments in the state of the art and comments received from users of the standard. Every IEEE Standard is subjected to review at least every five years for revision or reaffirmation, or every ten years for stabilization. When a document is more than five years old and has not been reaffirmed, or more than ten years old and has not been stabilized, it is reasonable to conclude that its contents, although still of some value, do not wholly reflect the present state of the art. Users are cautioned to check to determine that they have the latest edition of any IEEE Standard.

In publishing and making this document available, the IEEE is not suggesting or rendering professional or other services for, or on behalf of, any person or entity. Nor is the IEEE undertaking to perform any duty owed by any other person or entity to another. Any person utilizing this, and any other IEEE Standards document, should rely upon his or her independent judgment in the exercise of reasonable care in any given circumstances or, as appropriate, seek the advice of a competent professional in determining the appropriateness of a given IEEE standard.

Interpretations: Occasionally questions may arise regarding the meaning of portions of standards as they relate to specific applications. When the need for interpretations is brought to the attention of IEEE, the Institute will initiate action to prepare appropriate responses. Since IEEE Standards represent a consensus of concerned interests, it is important to ensure that any interpretation has also received the concurrence of a balance of interests. For this reason, IEEE and the members of its societies and Standards Coordinating Committees are not able to provide an instant response to interpretation requests except in those cases where the matter has previously received formal consideration. A statement, written or oral, that is not processed in accordance with the IEEE-SA Standards Board Operations Manual shall not be considered the official position of IEEE or any of its committees and shall not be considered to be, nor be relied upon as, a formal interpretation of the IEEE. At lectures, symposia, seminars, or educational courses, an individual presenting information on IEEE standards shall make it clear that his or her views should be considered the personal views of that individual rather than the formal position, explanation, or interpretation of the IEEE.

Comments for revision of IEEE Standards are welcome from any interested party, regardless of membership affiliation with IEEE. Suggestions for changes in documents should be in the form of a proposed change of text, together with appropriate supporting comments. Recommendations to change the status of a stabilized standard should include a rationale as to why a revision or withdrawal is required. Comments and recommendations on standards, and requests for interpretations should be addressed to:

Secretary, IEEE-SA Standards Board

445 Hoes Lane

Piscataway, NJ 08854

USA

Authorization to photocopy portions of any individual standard for internal or personal use is granted by The Institute of Electrical and Electronics Engineers, Inc., provided that the appropriate fee is paid to Copyright Clearance Center. To arrange for payment of licensing fee, please contact Copyright Clearance Center, Customer Service, 222 Rosewood Drive, Danvers, MA 01923 USA; +1 978 750 8400. Permission to photocopy portions of any individual standard for educational classroom use can also be obtained through the Copyright Clearance Center.

Introduction

This introduction is not part of IEEE P802.15.4j/D00, Draft Standard for Amendment: Alternative Physical Layer Extension to support Medical Body Area Network (MBAN) services operating in the 2360-2400 MHz band.

Notice to users

Laws and regulations

Users of these documents should consult all applicable laws and regulations. Compliance with the provisions of this standard does not imply compliance to any applicable regulatory requirements. Implementers of the standard are responsible for observing or referring to the applicable regulatory requirements. IEEE does not, by the publication of its standards, intend to urge action that is not in compliance with applicable laws, and these documents may not be construed as doing so.

Copyrights

This document is copyrighted by the IEEE. It is made available for a wide variety of both public and private uses. These include both use, by reference, in laws and regulations, and use in private self-regulation, standardization, and the promotion of engineering practices and methods. By making this document available for use and adoption by public authorities and private users, the IEEE does not waive any rights in copyright to this document.

Updating of IEEE documents

Users of IEEE standards should be aware that these documents may be superseded at any time by the issuance of new editions or may be amended from time to time through the issuance of amendments, corrigenda, or errata. An official IEEE document at any point in time consists of the current edition of the document together with any amendments, corrigenda, or errata then in effect. In order to determine whether a given document is the current edition and whether it has been amended through the issuance of amendments, corrigenda, or errata, visit the IEEE Standards Association web site at http://ieeexplore.ieee.org/xpl/standards.jsp, or contact the IEEE at the address listed previously.

For more information about the IEEE Standards Association or the IEEE standards development process, visit the IEEE-SA web site at http://standards.ieee.org.

Errata

Errata, if any, for this and all other standards can be accessed at the following URL: 
http://standards.ieee.org/reading/ieee/updates/errata/index.html. Users are encouraged to check this URL for errata periodically.

Interpretations

Current interpretations can be accessed at the following URL: http://standards.ieee.org/reading/ieee/interp/
index.html.

Patents

[If the IEEE has not received letters of assurance prior to the time of publication, the following notice shall appear:]

Attention is called to the possibility that implementation of this standard may require use of subject matter covered by patent rights. By publication of this standard, no position is taken with respect to the existence or validity of any patent rights in connection therewith. The IEEE is not responsible for identifying Essential Patent Claims for which a license may be required, for conducting inquiries into the legal validity or scope of Patents Claims or determining whether any licensing terms or conditions provided in connection with submission of a Letter of Assurance, if any, or in any licensing agreements are reasonable or non-discriminatory. Users of this standard are expressly advised that determination of the validity of any patent rights, and the risk of infringement of such rights, is entirely their own responsibility. Further information may be obtained from the IEEE Standards Association.

[The following notice shall appear when the IEEE receives assurance from a known patent holder or patent applicant prior to the time of publication that a license will be made available to all applicants either without compensation or under reasonable rates, terms, and conditions that are demonstrably free of any unfair discrimination.]

Attention is called to the possibility that implementation of this standard may require use of subject matter covered by patent rights. By publication of this standard, no position is taken with respect to the existence or validity of any patent rights in connection therewith. A patent holder or patent applicant has filed a statement of assurance that it will grant licenses under these rights without compensation or under reasonable rates, with reasonable terms and conditions that are demonstrably free of any unfair discrimination to applicants desiring to obtain such licenses. Other Essential Patent Claims may exist for which a statement of assurance has not been received. The IEEE is not responsible for identifying Essential Patent Claims for which a license may be required, for conducting inquiries into the legal validity or scope of Patents Claims, or determining whether any licensing terms or conditions provided in connection with submission of a Letter of Assurance, if any, or in any licensing agreements are reasonable or non-discriminatory. Users of this standard are expressly advised that determination of the validity of any patent rights, and the risk of infringement of such rights, is entirely their own responsibility. Further information may be obtained from the IEEE Standards Association.

Participants

At the time this draft Amendment was submitted to the IEEE-SA Standards Board for approval, the IEEE P802.15 Working Group had the following membership:

Robert F Heile, Chair

Rick Alfin, Co-Vice Chair

Patrick W. Kinney, Co-Vice Chair and Secretary

James P. K. Gilb, Working Group Technical Editor

Clint Chaplin, Treasurer

Raymond Krasinski, Task Group 4j Chair

David H. Evans, Task Group 4j Technical Editor and Secretary

The following members of the IEEE 802.15 balloting committee voted on this draft. Balloters may have voted for approval, disapproval, or abstention.

(to be supplied by IEEE)

Balloter1

Balloter2

Balloter3

Balloter4

Balloter5

Balloter6

Balloter7

Balloter8

Balloter9

Major contributions besides the chair, secretary and the editor were received from the following individuals:

Anuj Batra

Paul Chilton

Jin-Meng Ho

Suhwook Kim

Seung-Hoon Park

Ranjeet K. Patro

Jaeseung Son

Dong Wang

Ariton Xhafa

Betty Zhao

The following members of the balloting committee voted on this draft. Balloters may have voted for approval, disapproval, or abstention.

(to be supplied by IEEE)

Balloter1

Balloter2

Balloter3

Balloter4

Balloter5

Balloter6

Balloter7

Balloter8

Balloter9

Contents

1. Overview	1

2. Normative references	1

3. Definitions	2

3.1 Definitions	2

3.2 Acronyms and abbreviations	2

4. General description	2

5. MAC protocol	2

5.1 MAC functional description	2

5.1.1 Channel access	2

5.1.2 Starting and maintaining PANs	2

5.1.3 Association and disassocisation	6

5.1.4 Synchronization	8

5.1.5 Transaction handling	8

5.1.6 Transmission, reception, and acknowledgement	8

5.1.7 GTS allocation and management	8

5.2 MAC frame format	10

5.2.1 General MAC frame format	10

5.2.2 Format of individual frame types	10

5.3 MAC command frames	11

5.3.1 Association request command	12

5.3.2 Association response command	12

5.3.3 Disassociation notification command	12

5.3.4 Data request command	12

5.3.5 PAN ID conflict notification command	12

5.3.6 Orphan notification command	12

5.3.7 Beacon request command	12

5.3.8 Coordinator realignment command	12

5.3.9 GTS request command	12

5.3.10 Channel switch notification command	14

5.3.11 Grant association proxy request	15

5.3.12 Grant association proxy response	16

5.3.13 Association proxy request	17

5.3.14 Association proxy response	18

5.3.15 Coordinator switch	19

6. MAC services	21

6.1 Overview	21

6.2 MAC management service	21

6.2.1 Common requirements for MLME primitives	21

6.2.2 Association primitives	21

6.2.3 Disassociation primitives	22

6.2.4 Communications notification primitives	22

6.2.5 Primitives for reading PIB attributes	22

6.2.6 GTS management primitives	22

6.2.7 Primitives for orphan notification	22

6.2.8 Primitives for resetting the MAC sublayer	22

6.2.9 Primitives for specifying the receiver enable time	22

6.2.10 Primitives for channel scanning	22

6.2.11 Primitives for writing PIB attributes	22

6.2.12 Primitives for updating the superframe configuration	22

6.2.13 Primitives for synchronizing with a coordinator	22

6.2.14 Primitives for requesting data from a coordinator	22

6.2.15 Primitives for specifying dynamic preamble	22

6.2.16 Primitives for channel sounding	22

6.2.17 Primitives for ranging calibration (for UWB PHYs)	22

6.2.18 Channel switch notification primitives	22

6.2.19 Periodic GTS management primitives	26

6.2.20 Grant association proxy notification primitives	28

6.2.21 Grant association proxy notification primitives	32

6.2.22 Coordinator switch notification primitives	35

6.3 MAC data service	39

6.4 MAC constants and PIB attributes	39

6.4.1 MAC constants	39

6.4.2 MAC PIB attributes	39

7. Security	40

8. General PHY requirements	40

8.1 General requirements and definitions	40

8.1.1 Operating frequency range	40

8.1.2 Channel assignments	40

9. PHY services	41

10. O-QPSK PHY	41

10.1 PPDU format	41

10.2 Modulation and spreading	41

10.2.1 Data rate	41

10.2.2 Reference modulator diagram	41

10.2.3 Bit-to-symbol mapping	41

10.2.4 Symbol-to-chip mapping	41

10.2.5 O-QPSK modulation	42

10.2.6 Pulse shape	42

10.2.7 Chip transmission order	42

10.3 O-QPSK PHY RF requirements	42

10.3.1 Operating frequency range	42

10.3.2 Transmit power spectral density (PSD) mask	43

10.3.3 Symbol rate	43

10.3.4 Receiver sensitivity	43

10.3.5 Receiver interference rejection	43

Annex I	44

Channel bitmap	44

IEEE Standard for

Local and metropolitan area networks–

Part 15.4: Low-Rate Wireless Personal Area Networks (LR-WPANs)

Amendment: Alternative Physical Layer Extension to support Medical Body Area Networks (MBANS) services operating in the 2360-2400 MHz band

NOTE—The editing instructions contained in this amendment define how to merge the material contained therein into the existing base standard and its amendments to form the comprehensive standard.

The editing instructions are shown in bold italic. Four editing instructions are used: change, delete, insert, and replace. Change is used to make corrections in existing text or tables. The editing instruction specifies the location of the change and describes what is being changed by using strikethrough (to remove old material) and underscore (to add new material). Delete removes existing material. Insert adds new material without disturbing the existing material. Deletions and insertions may require renumbering. If so, renumbering instructions are given in the editing instruction. Replace is used to make changes in figures or equations by removing the existing figure or equation and replacing it with a new one. Editing instructions, change markings, and this NOTE will not be carried over into future editions because the changes will be incorporated into the base standard.

Overview

Normative references

The following referenced documents are indispensable for the application of this document (i.e., they must be understood and used, so each referenced document is cited in text and its relationship to this document is explained). For dated references, only the edition cited applies. For undated references, the latest edition of the referenced document (including any amendments or corrigenda) applies.

Insert the following new reference alphabetically into Clause 2:

IEEE Std 802.1D-2004, IEEE Standard for Local and Metropolitan Area Networks: Media Access Control (MAC) Bridges.

Definitions

Definitions

Acronyms and abbreviations

Insert the following acronym alphabetically into 3.2:

MBANS		medical body area network services

General description

MAC protocol

MAC functional description

Channel access

Starting and maintaining PANs

Scanning through channels

PAN identifier conflict resolution

Starting and realigning a PAN

Starting a PAN

Realigning a PAN

Realignment in a PAN

Updating superframe configuration and channel PIB attributes

Insert the following after 5.1.2.3.4:

Channel switch notification

The channel switch notification procedure is initiated by the next higher layer by issuing the MLME-CHANNELSWITCH.request primitive, as described in 6.2.18.1, to the MLME.

When a coordinator wants one of its associated devices to switch its operating channel at a specific time, the MLME of the coordinator shall send the channel switch notification command in the manner specified by the TxIndirect parameter of the MLME-CHANNELSWITCH.request primitive previously sent by the next higher layer. If TxIndirect is TRUE, the MLME of the coordinator shall send the channel switch notification command to the device using indirect transmission; i.e. the channel switch notification command frame shall be added to the list of pending transactions stored on the coordinator and extracted at the discretion of the device concerned using the method described in 5.1.6.3. If the command frame is not successfully extracted by the device, the coordinator shall consider the device disassociated. Otherwise, the MLME shall send the channel switch notification command to the device directly. In this case, if the channel switch notification command cannot be sent due to a channel access failure, the MAC sublayer shall notify the next higher layer.

Upon successful transmission of the channel switch command, the MAC sublayer shall issue the MLME-CHANNELSWITCH.confirm primitive with a status of SUCCESS.

If a device has received the channel switch command, as defined in 5.3.10, from the coordinator, the MLME shall issue the MLME-CHANNELSWITCH.indication primitive with the NewPANID, CoordinatorAddress, RemainingTime, ChannelNumber and ChannelPage parameters set to the respective field in the channel switch command.

Figure 16a illustrates the sequence necessary for a coordinator in a beacon-enabled PAN to successfully transmit channel switch notification command to a device using indirect transmission.

Figure 16a—Message sequence chart for channel switch notification initiated by a coordinator, using indirect transmission, in a beacon-enabled PAN

Beacon generation

Device discovery

Insert the following after 5.1.2.5:

Coordinator switch

A PAN coordinator can switch the devices that are associated to its PAN to another coordinator by using the coordinator switch and device re-alignment commands. The PAN coordinator shall first broadcast a coordinator switch request on a channel using the coordinator switch request command described in 5.3.15 indicating the number of devices that it wishes to switch. The number of devices defines the devices that are associated with its PAN. If another coordinator has the resources and is able to accept the number of devices, it may respond to the broadcast coordinator switch request by sending a coordinator switch response command to the PAN coordinator confirming that it is able to accept the defined number of devices. The time for which the PAN coordinator listens for responses after transmitting the coordinator request command is beyond the scope of this standard. The PAN coordinator shall repeat the broadcast of the coordinator switch request on all of the available channels. The sequence of coordinator switch broadcast request and coordinator switch response messages are shown in Figure 16b.

Figure 16b – Coordinator switch broadcast and response

The PAN coordinator shall then select a new coordinator and confirm to the new coordinator using the coordinator switch request command that the new coordinator shall accept the association of the number of devices defined in the command. The process for the selection of a new coordinator is beyond the scope of this standard. Once the PAN coordinator has received a successful confirmation of the coordinator switching request command from the new coordinator, it shall send a channel switch notification command to its devices containing the PAN ID and the extended address of the new coordinator and the channel number and channel page that it is operating on. On receiving the channel switch notification command, the devices shall attempt to associate with the new coordinator whose details are provided in this command and shall do so without performing a channel scan. The sequence of coordinator switch selection and channel switch notification messages are shown in Figure 16c.

Figure 16c – Coordinator switch selection and channel switch notification

Association and disassociation

Association

Disassociation

Insert the following after5.1.3.2:

Association proxy

A device shall attempt a request for grant of association proxy after having first associated with a suitable PAN. A coordinator shall allow a grant of association proxy only if macAssociationPermit is set to TRUE. Similarly, a device shall attempt a request for grant of association proxy only if the coordinator is currently allowing association, as indicated in the results of the scanning procedure. If a coordinator with macAssociationPermit set to FALSE receives a grant association proxy request command from a device, the coordinator shall ignore the command.

The MAC sublayer shall initiate the grant association proxy procedure by sending a grant association proxy request command, as described in 5.3.11, to the coordinator of the PAN; if the grant association proxy request command cannot be sent due to a channel access failure, the MAC sublayer shall notify the next higher layer.

The acknowledgment to a grant association proxy request command does not mean that the device has associated. The coordinator needs time to determine whether the current resources available on the PAN are sufficient to allow other devices to associate. The coordinator shall make this decision within macResponseWaitTime of receiving this command. If the coordinator has sufficient resources, it shall allocate a set of short addresses to the devices, and its MAC sublayer shall generate a grant association proxy response command, as described in 5.3.12. The command shall specify the number of devices that have been allocated a short address, the actual short addresses that the devices may use and a status indicating a successful grant of proxy association. If the coordinator has insufficient resources its MLME shall generate a grant association proxy response command containing a status indicating a failure, as defined in Table 6. The grant association proxy response command shall be sent to the device requesting the grant of association proxy using indirect transmission; i.e., the grant association proxy response command frame shall be added to the list of pending transactions stored on the coordinator and extracted at the discretion of the device concerned using the method described in 5.1.6.3.

On receipt of the acknowledgment to the grant association proxy request command, the device shall wait for at most macResponseWaitTime from when the command is received for the coordinator to make its grant of association proxy decision; the PIB attribute macResponseWaitTime is a network-topology-dependent parameter and shall be set to match the specific requirements of the network that a device is trying to join. If the device is tracking the beacon, it shall attempt to extract the grant association proxy response command from the coordinator whenever it is indicated in the beacon frame. If the device is not tracking the beacon, it shall attempt to extract the grant association proxy response command from the coordinator after macResponseWaitTime. If the device does not extract a grant association proxy response command frame from the coordinator within macResponseWaitTime, the MLME shall issue the MLME-GRANTASSOCIATIONPROXY.confirm primitive, as described in 6.2.20.4, with a status of NO_DATA, and the request for grant of association proxy attempt shall be deemed a failure. In this case, the device shall terminate any tracking of the beacon, by issuing an MLME-SYNC.request primitive, as described in 6.2.13.1, with the TrackBeacon parameter set to FALSE.

A device shall attempt a request for an association proxy after obtaining the extended address and capability information for each of the devices for which it has requested a grant of association proxy. The method for obtaining this information from the devices is outside the scope of this standard.

The MAC sublayer shall initiate the association proxy procedure by sending an association proxy request command, as described in 5.3.13, to the coordinator of the PAN.

On receipt of an association proxy request command, the coordinator shall update the device-specific information for the allocated short address with the device extended address and the device capability information. If the coordinator finds that the device was previously associated with its PAN, it shall replace all previously obtained device-specific information. If the coordinator has successfully updated the device information, it shall respond with an association proxy response command indicating success.

A message sequence chart for grant of an association proxy is illustrated in Figure 19a.

Figure 19a—Message sequence chart for grant of an association proxy

A message sequence chart for an association proxy is illustrated in Figure 19b.

Figure 19b—Message sequence chart for an association proxy

Synchronization

Transaction handling

Transmission, reception, and acknowledgement

GTS allocation and management

CAP maintenance

GTS allocation

Insert after the second paragraph:

To request the allocation of a new periodic GTS, the MLME shall send the periodic GTS request command, as described in 5.3.9, to the PAN coordinator. The Characteristics Type field of the request shall be set to one, and the length, direction, start frame and period exponent fields shall be set according to the desired characteristics of the required periodic GTS.

Insert after the fifth paragraph:

When the PAN coordinator determines whether capacity is available for the requested periodic GTS, it shall generate a periodic GTS descriptor with the requested specifications and the short address of the requesting device. If the periodic GTS is allocated successfully, the PAN coordinator shall set the start slot in the periodic GTS descriptor to the superframe slot at which the periodic GTS begins and the GTS BSN in the periodic GTS descriptor to the number of remaining superframes for the first allocated periodic GTS slot. In addition, the PAN coordinator shall notify the next higher layer of the new periodic GTS. This notification is achieved when the MLME of the PAN coordinator issues the MLME-PERIODIC-GTS.indication primitive, as described in 6.2.19.3 with the characteristics of the allocated periodic GTS.  If there is not sufficient capacity to allocate the requested GTS, the start slot and GTS BSN shall be set to zero. The PAN coordinator shall then include this periodic GTS descriptor in its beacon and update the GTS Specification field of the beacon frame accordingly. The PAN coordinator shall also update the Final CAP Slot field of the Superframe Specification field of the beacon frame, indicating the final superframe slot utilized by the decreased CAP. The periodic GTS descriptor shall remain in the beacon frame for aGTSDescPersistenceTime superframes, after which it shall be removed automatically. The PAN coordinator shall be allowed to reduce its CAP below aMinCAPLength to accommodate the temporary increase in the beacon frame length due to the inclusion of the periodic GTS descriptor.

GTS usage

GTS deallocation

Insert after the first paragraph:

A device is instructed to request the deallocation of an existing periodic GTS through the MLME-PERIODIC-GTS.request primitive, as described in 6.2.19.1, using the characteristics of the periodic GTS it wishes to deallocate. From this point onward, the periodic GTS to be deallocated shall not be used by the device, and its stored characteristics shall be reset.

Change and insert in the second paragraph as shown:

To request the deallocation of an existing GTS, the MLME shall send the GTS request command, as described in 5.3.9, to the PAN coordinator. To request the deallocation of an existing periodic GTS, the MLME shall send the periodic GTS request command, as described in 5.3.9, to the PAN coordinator. The Characteristics Type field of either the GTS Characteristics field or the Periodic GTS Characteristics field of the request shall be set to zero (i.e., GTS deallocation), and the length and direction fields shall be set according to the characteristics of the GTS to deallocate. On receipt of the acknowledgment to either the GTS request command or the periodic GTS request command, the MLME shall notify the next higher layer of the deallocation. This notification is achieved when either the MLME issues the MLME-GTS.confirm primitive, as described in 6.2.6.2, with a status of SUCCESS and a GTSCharacteristics parameter with its Characteristics Type field set to zero or when the MLME issues the MLME-PERIODIC-GTS.confirm primitive, as described in 6.2.19.2, with a status of SUCCESS and a PeriodicGTSCharacteristics parameter with its Characteristics Type field set to zero. If either the GTS request command or the periodic GTS request command is not received correctly by the PAN coordinator, it shall determine that the device has stopped using either its GTS or its periodic GTS by the procedure described in 5.1.7.6.

GTS reallocation

GTS expiration

Insert at the end of 5.1.7.6:

The MLME of the PAN coordinator shall attempt to detect when a device has stopped using a periodic GTS using the following rules:

For a transmit periodic GTS, the MLME of a PAN coordinator shall assume that a device is no longer using its periodic GTS if a data frame is not received from the device in the GTS at least every 2 × m superframes, where m is defined below.

For a receive periodic GTS, the MLME of the PAN coordinator shall assume that a device is no longer using its periodic GTS if an acknowledgment frame is not received from the device at least every 2 × m superframes. If the data frames sent in the GTS do not require acknowledgment frames, the MLME of the PAN coordinator will not be able to detect whether a device is using its receive periodic GTS. However, the PAN coordinator is capable of deallocating the periodic GTS at any time.

The value of m is defined as follows:

m = P * 2(8-macBeaconOrder)	0 ≤ macBeaconOrder ≤ 8

m = P		9 ≤ macBeaconOrder ≤ 14

where P is defined in 5.3.9.3.

MAC frame format

General MAC frame format

Format of individual frame types

Beacon frame format

Beacon frame MHR fields

Superframe Specification field

GTS Specification field

Change and insert in figure 42 as shown:

| Bits: 0-2 | 3-5 | 6 | 6 |
| --- | --- | --- | --- |
| GTS Descriptor Count | Reserved | Periodic GTS Permit | GTS Permit |

Figure 42–Format of the GTS Specification field

Insert at the end of 5.2.2.1.3:

The Periodic GTS Permit field shall be set to one if macPeriodicGTSPermit is equal to TRUE (i.e., the PAN coordinator is accepting periodic GTS requests). Otherwise, the Periodic GTS Permit field shall be set to zero.

GTS Directions field

GTS List field

Insert at the end of 5.2.2.1.5:

If a GTS descriptor represents a periodic GTS (i.e. the GTS descriptor corresponds to a periodic GTS request), the GTS Length field shall contain the four least significant bits of the macBSN indicating the first superframe that the periodic GTS is available for the addressed device.

MAC command frames

Insert and change as shown in Table 5:

Table 5–MAC command frames

| Command frame identifier | Command frame | Tx | Rx | Subclause |
| --- | --- | --- | --- | --- |
| 0x0a | Channel switch notification |  |  | 5.3.10 |
| 0x0b | Grant association proxy request |  |  | 5.3.11 |
| 0x0c | Grant association proxy response |  |  | 5.3.12 |
| 0x0d | Association proxy request |  |  | 5.3.13 |
| 0x0e | Association proxy response |  |  | 5.3.14 |
| 0x0f | Coordinator switch request |  |  | 5.3.15.1 |
| 0x1a | Coordinator switch response |  |  | 5.3.15.2 |
| 0x1b-0xff | Reserved |  |  | - |

Association request command

Association response command

MHR fields

Short Address field

Association Status field

Insert and change as shown in Table 6:

Table 6–Valid values of the Association Status field

| Association Status | Description |
| --- | --- |
| 0x03-0x7f0x0f | Reserved |
| 0xa0-0xbf | Number of allocated device addresses where the number is offset by 0xa0. |
| 0xc0-0x7f | Reserved. |

Disassociation notification command

Data request command

PAN ID conflict notification command

Orphan notification command

Beacon request command

Coordinator realignment command

GTS request command

Insert as shown in the first paragraph of 5.3.9 as follows:

The GTS request command is used by an associated device that is requesting the allocation of either a new GTS or a new periodic GTS or the deallocation of either an existing GTS or an existing periodic GTS from the PAN coordinator. Only devices that have a short address less than 0xfffe shall send this command.

Change and insert in Figure 58:

| Octets: 7 | 1 | 1/2 |
| --- | --- | --- |
| MHR fields | Command Frame Identifier | GTS Characteristics/ Periodic GTS Characteristics |

Figure 58–GTS request command format

MHR fields

GTS Characteristics field

Insert after 5.3.9.2:

Periodic GTS Characteristics field

When the GTS request command is used by an associated device that is requesting the allocation of a new periodic GTS or the deallocation of an existing periodic GTS from the PAN coordinator, the Periodic GTS Characteristics field shall be included.

The Periodic GTS Characteristics field shall be formatted as illustrated in Figure 59a.

| Bits: 0-3 | 4 | 5 | 6-7 | 8-11 | 12-14 | 15 |
| --- | --- | --- | --- | --- | --- | --- |
| GTS Length | GTS Direction | Characteristic Type | Reserved | Start Frame | GTS Period Exponent | Reserved |

Figure 59a–Periodic GTS Characteristic field format

The GTS Length field shall contain the number of superframe slots being requested for the periodic GTS.

The GTS Direction field shall be set to one if the periodic GTS is to be a receive-only periodic GTS. Conversely, this field shall be set to zero if the GTS is to be a transmit-only periodic GTS. GTS direction is defined relative to the direction of data frame transmissions by the device.

The Characteristics Type field shall be set to one if the characteristics refer to a periodic GTS allocation or zero if the characteristics refer to a periodic GTS deallocation.

The Start Frame field shall contain the value S, 0 ≤ S ≤ 7, which indicates the first periodic GTS for the requesting device is being requested to occur in a superframe no later than (S+1) superframes after the current superframe.

The GTS Period Exponent field shall specify the value, N, and this shall be used to define the period P, in superframes that is being requested for the periodic GTS.

The value of P is defined as follows:

P = 2(N+1)	for 0 ≤ N ≤ 7

Insert after 5.3.9:

Channel switch notification command

The channel switch notification command is used by a device to notify another device to switch a channel at certain time. In MBANS, this command shall only be sent by a PAN coordinator. All devices shall be capable of receiving this command.

This command is optional.

The channel switch notification command shall be formatted as illustrated in Figure 59b.

| Octets: variable | 1 | variable | 1 | 1 |
| --- | --- | --- | --- | --- |
| MHR fields | Command Frame Identifier | Standard Dependent Information | Channel Number | Channel Page |

Figure 59b–Channel switch notification command format

MHR fields

The Destination Addressing Mode field and the Source Addressing Mode field shall be set to indicate extended addressing for MBANS devices.

The Frame Pending field shall be set to zero and ignored upon reception. The AR field shall be set to one. The Frame Version field shall be set as specified in 5.2.3.

For MBANS devices, the Destination PAN Identifier field shall contain the broadcast PAN identifier. The Destination Address field shall contain the extended address of the destination device. The Source PAN Identifier field shall contain the value of macPANId, and the Source Address field shall contain the value of macExtendedAddress.

Standard Dependent Information

MBANS information

The MBANS information is formatted as illustrated in figure 59c.

| Octets: 2 | 2/8 | 2 |
| --- | --- | --- |
| New PAN ID | Coordinator Address | Remaining Time |

Figure 59c–MBANS information format

New PAN ID field

The New PAN ID field shall contain the PAN ID of the PAN coordinator that operates in the channel specified in the Channel Number field and Channel Page field. The receiving device is requested to associate with the PAN coordinator for the following communication after channel switching.

Coordinator Address field

The Coordinator Address field shall contain the coordinator address of the PAN coordinator that operates in the channel specified in the Channel Number field and Channel Page field. The coordinator address field shall contain either the coordinator short address or coordinator extended address.

Remaining Time field

The Remaining Time field shall contain the remaining time in minutes after which the channel switching shall occur.

Channel Number field

The Channel Number field shall contain the channel number that the initiating device intends to use for the following communication, and the receiving device is requested to switch to this channel number.

Channel Page field

The Channel Page field shall contain the channel page that the initiating device intends to use for the following communication, and the receiving device is requested to switch to one channel of this channel page.

Insert:

Grant association proxy request

The grant association proxy request command allows a FFD to associate a number of RFDs to a PAN coordinator.

Only FFDs shall be capable of transmitting and receiving this command. The FFD shall be associated with the PAN coordinator before using the grant association proxy request command.

This command is optional.

The grant association proxy request command shall be formatted as illustrated in Figure 59d.

| Octets: variable | 1 | 1 |
| --- | --- | --- |
| MHR fields | Command Frame Identifier | Device Number |

Figure 59d – Grant association proxy request command format

MHR fields

The Destination Addressing Mode field shall be set to indicate long addressing. The Source Addressing Mode field shall be set to indicate long addressing.

The Frame Pending field shall be set to zero and ignored upon reception and the AR field shall be set to one.

The Destination PAN Identifier field shall contain the identifier of the PAN to which the grant association proxy is requested. The Destination Address field shall contain the address from the beacon frame that is transmitted by the coordinator to which the grant associated proxy request command is being sent. The Source PAN Identifier field shall contain the broadcast PAN identifier. The Source Address field shall contain the value of macExtendedAddress.

Device Number fields

The Device Number field shall be formatted as illustrated in Figure 59e.

| Bits: 0-4 | 5-7 |
| --- | --- |
| Number of Devices | Reserved |

Figure 59e–Device Number field format

The Number of Devices shall be set to indicate the total number of devices for which device short addresses are being requested.

Grant association proxy response

The grant association proxy response command allows a PAN coordinator to communicate the results of a grant association proxy attempt back to the device requesting a grant association proxy.

This command shall only be sent by the PAN coordinator to a FFD that is currently trying to request a grant association proxy.

Only FFDs shall be capable of receiving this command.

The grant association proxy response command shall be formatted as illustrated in Figure 59f.

| Octets: variable | 1 | 1 | 2*A | 1 |
| --- | --- | --- | --- | --- |
| MHR fields | Command Frame Identifier | Number of Allocated Short Addresses (A) | Short Address 1 | … | Short Address A | Association Status |

Figure 59f–Grant association proxy response command format

MHR fields

The Destination Addressing Mode and the Source Addressing Mode fields shall each be set to indicate extended addressing.

The Frame Pending field shall be set to zero and ignored upon reception and the AR field shall be set to one.

The PAN ID Compression field shall be set to one. In accordance with this value of the PAN ID Compression field, the Destination PAN ID field shall contain the value of macPANId, while the Source PAN ID field shall be omitted.

The Destination Address field shall contain the extended address of the device requesting association. The Source Address field shall contain the value of macExtendedAddress.

Number of Allocated Short Addresses field

If the coordinator is able to associate a number of devices to its PAN, this field shall contain the number, A,  of short addresses that the coordinator is able to allocate, where 1 ≤ A ≤ 32.

Short Address fields

The number of short address fields shall be equal to A, the number of short addresses that the coordinator was able to allocate. These fields shall contain the short addresses that the devices shall use in their communications on the PAN until they are disassociated.

Association Status field

Valid values of the Association Status field are defined in 5.3.2.3 and Table 6.

Association proxy request

The association proxy request command allows a FFD to update a PAN coordinator with the extended addresses and the Capability Information of the RFDs that have been associated using the grant association proxy request and response commands.

Only FFDs shall be capable of transmitting and receiving this command.

This command is optional.

The association proxy request command shall be formatted as illustrated in Figure 59g.

| Octets: variable | 1 | 2 | 8 | 1 |
| --- | --- | --- | --- | --- |
| MHR fields | Command Frame Identifier | Device Short Address | Device Extended Address | Capability Information |

Figure 59g–Association proxy request command format

MHR fields

The Destination Addressing Mode field shall be set to indicate long addressing. The Source Addressing Mode field shall be set to indicate long addressing.

The Frame Pending field shall be set to zero and ignored upon reception and the AR field shall be set to one.

The Destination PAN Identifier field shall contain the identifier of the PAN to which the association proxy is requested.

Device Short Address

The Device Short Address field shall contain the short address of the device whose extended address and Capability Information is being transferred to the PAN coordinator.

Device Extended Address

The Device Extended Address field shall contain the extended address of the device whose information is being updated in the PAN coordinator.

Capability Information field

The Capability Information field shall contain the information of the device whose information is being updated in the PAN coordinator, and shall be formatted as illustrated in figure 50.

Association proxy response

The association proxy response command allows a PAN coordinator to communicate the results of an association proxy attempt to the device requesting an association proxy.

This command shall only be sent by the PAN coordinator to a FFD that is currently trying to request an association proxy.

Only FFDs shall be capable of transmitting and receiving this command.

The association proxy response command shall be formatted as illustrated in Figure 59h.

| Octets: variable | 1 | 2 | 1 |
| --- | --- | --- | --- |
| MHR fields | Command Frame Identify | Short Address | Association Status |

Figure 59h–Association proxy response command format

MHR fields

The Destination Addressing Mode field shall be set to indicate long addressing. The Source Addressing Mode field shall be set to indicate long addressing.

The Frame Pending field shall be set to zero and ignored upon reception and the AR field shall be set to one.

The PAN ID Compression field shall be set to one. In accordance with this value of the PAN ID Compression field, the Destination PAN ID field shall contain the value of macPANId, while the Source PAN ID field shall be omitted.

The Destination Address field shall contain the extended address of the device requesting association. The Source Address field shall contain the value of macExtendedAddress.

Short Address field

If the coordinator is not able to associate the device to its PAN, the Short Address field shall be set to 0xffff, and the Association Status field shall contain the reason for the failure.

If the coordinator is able to update the extended address and the Capability Information of the device that had been associated using the association proxy request, this field shall contain the short address of that device.

Association Status field

Valid values of the Association Status field are defined in 5.3.2.3 and in Table 6.

Coordinator switch

Coordinator switch request

The coordinator switch request command allows a PAN coordinator to determine whether there are any other coordinators that have the resources to accept the devices that are associated with the PAN coordinator. The coordinator switch request command also allows a PAN coordinator to confirm to a new coordinator that it has selected that coordinator to accept the devices that are associated with its PAN.

Only FFDs shall be capable of transmitting and receiving this command.

This command is optional.

The coordinator switch request command shall be formatted as illustrated in Figure 59i.

| Octets: variable | 1 | 1 |
| --- | --- | --- |
| MHR fields | Command Frame Identifier | Number of Devices |

Figure 59i–Coordinator switch request command format

MHR fields

The Destination Addressing Mode field shall be set to indicate short addressing when the coordinator switch request is broadcast in a channel. Otherwise the Destination Addressing Mode field shall be set to indicate extended addressing. The Source Addressing Mode field shall be set to indicate extended addressing.

The Frame Pending field shall be set to zero and ignored upon reception and the AR field shall be set to zero.

The Destination PAN Identifier field shall contain the broadcast PAN identifier and the Destination Address field shall contain the broadcast short address when the coordinator switch request is broadcast in a channel. Otherwise the Destination PAN Identifier field shall contain the macPANID of the new PAN coordinator and the Destination Address field shall contain the macExtendedAddress of the new PAN coordinator when the coordinator switch request is transmitted to a new PAN coordinator.

The Source PAN Identifier field shall contain the value of the macPANID of the source PAN coordinator. The Source Address field shall contain the value of macExtendedAddress of the source PAN coordinator.

Number of Devices field

The number of device field shall be formatted as illustrated in Figure 59j.

Bits: 0-7

Number of Devices

Figure 59j–Number of Devices field format

The Number of Devices shall be used to indicate the total number of devices that are associated with the source PAN coordinator.

Coordinator switch response

The coordinator switch response command allows another coordinator to communicate the results of a coordinator switch request attempt back to the source PAN coordinator requesting a coordinator switch.

This command shall only be sent by FFDs to the source PAN coordinator that is currently trying to request a coordinator switch.

Only FFDs shall be capable of receiving this command.

The coordinator switch response command shall be formatted as illustrated in Figure 59k.

| Octets: variable | 1 | 1 | 2 |
| --- | --- | --- | --- |
| MHR fields | Command Frame Identifier | Switch Status | New PAN ID |

Figure 59k–Coordinator switch response command format

MHR fields

The Destination Addressing Mode and the Source Addressing Mode fields shall each be set to indicate extended addressing.

The Frame Pending field shall be set to zero and ignored upon reception. If the coordinator switch response command is sent in response to a coordinator switch request command that was broadcast to the channel, the AR field shall be set to zero. Otherwise if the coordinator switch request command is sent directly to the new coordinator the AR field shall be set to one.

The PAN ID Compression field shall be set to zero. In accordance with this value of the PAN ID Compression field, the Destination PAN ID field shall contain the value of macPANId of the source PAN coordinator while the Source PAN ID field shall contain the broadcast PAN ID. The Destination Address field shall contain the extended address of the source PAN coordinator that sent the coordinator switch request command. The Source Address field shall contain the value of macExtendedAddress of the new coordinator that is responding to the PAN coordinator switch request.

Switch Status field

If the new coordinator is able to associate a number of devices to its PAN, this field shall contain the value of the Number of Devices field that is received on reception of the broadcast coordinator switch request command.

If the coordinator switch request command is sent directly to the new coordinator, it shall confirm that it is able to associate a number of devices to its PAN and this field shall contain the value of the Number of Devices field that is received on reception of the coordinator switch request command. Otherwise if the new coordinator is not able to associate a number of devices to its PAN, this field shall contain 0x00.

New PAN ID fields

The New PAN ID field shall contain the macPANId of the PAN coordinator that is responding to the coordinator switch request command.

MAC services

Overview

MAC management service

Insert as shown in Table 8:

Table 8 – Summary of the primitives accessed through the MLME-SAP

| Name | Request | Indication | Response | Confirm |
| --- | --- | --- | --- | --- |
| MLME-CHANNELSWITCH | 6.2.18.1* | 6.2.18.3* |  | 6.2.18.2* |
| MLME-PERIODIC-GTS | 6.2.19.1* | 6.2.19.3* |  | 6.2.19.2* |
| MLME-GRANTASSOCIATIONPROXY | 6.2.20.1* | 6.2.20.2* | 6.2.20.3* | 6.2.20.4* |
| MLME-ASSOCIATION | 6.2.21.1* | 6.2.21.2* |  | 6.2.21.3* |
| MLME-COORDINATOR-SWITCH | 6.2.22.1* | 6.2.22.2* | 6.2.22.3* | 6.2.22.4* |

Common requirements for MLME primitives

Association primitives

Disassociation primitives

Communications notification primitives

Primitives for reading PIB attributes

GTS management primitives

Primitives for orphan notification

Primitives for resetting the MAC sublayer

Primitives for specifying the receiver enable time

Primitives for channel scanning

Primitives for writing PIB attributes

Primitives for updating the superframe configuration

Primitives for synchronizing with a coordinator

Primitives for requesting data from a coordinator

Primitives for specifying dynamic preamble

Primitives for channel sounding

Primitives for ranging calibration (for UWB PHYs)

Insert after 6.2.17:

Channel switch notification primitives

These primitives are used by a device to switch operating channel at a specific time and to instruct another device to switch operating channel at a specific time.

MLME-CHANNELSWITCH.request

The MLME-CHANNELSWITCH.request primitive is used by a device to instruct another device to switch operating channel at a specific time. On receipt of the MLME-CHANNELSWITCH.request primitive, the MLME of the initiating device generates a channel switch notification command.

The semantics of this primitive are:

MLME-CHANNELSWITCH.request	(

DeviceAddrMode,

DeviceAddress,

ChannelNumber,

ChannelPage,

TxIndirect,

NewPANID,

CoordinatorAddress,

RemainingTime,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44a.

Table 44a–MLME-CHANNELSWITCH.request parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| DeviceAddrMode | Enumeration | SHORT_ADDRESS, EXTENDED_ADDRESS | The addressing mode of the device which is instructed to change its operating channel at a specific time. For MBANS devices, the default value is EXTENDED_ADDRESS |
| DeviceAddress | Device address | As specified by the DeviceAddrMode parameter | The address of the device which should change its operating channel at specific time. |
| Channel Number | Integer | Any valid channel number | As defined in 5.3.10.3. |
| ChannelPage | Integer | Any valid channel page | As defined in 5.3.10.4. |
| TxIndirect | Boolean | TRUE, FALSE | TRUE if the channel switch notification command is to be sent indirectly. |
| NewPANID | Integer | As defined in 5.3.10.2.1.1 | As defined in 5.3.10.2.1.1. |
| CoordinatorAddress | Device address | As defined in 5.3.10.2.1.2 | As defined in 5.3.10.2.1.2. |
| RemainingTime | Integer | As defined in 5.3.10.2.1.3 | As defined in 5.3.10.2.1.3. |
| SecurityLevel | Integer | As defined in Table 46 | As defined in Table 46. |
| KeyIdMode | Integer | As defined in Table 46 | As defined in Table 46. |
| KeySource | Set of octets | As defined in Table 46 | As defined in Table 46. |
| KeyIndex | Integer | As defined in Table 46 | As defined in Table 46. |

For MBANS devices, if this primitive is received by the MLME of a coordinator with the TxIndirect parameter set to TRUE, the channel switch notification command will be sent using indirect transmission as described in 5.1.5.

If this primitive is received by the MLME of a coordinator with the TxIndirect parameter set to FALSE, the MLME sends a channel switch notification command to the device in the CAP for beacon-enabled PAN.

MLME-CHANNELSWITCH.confirm

The MLME-CHANNELSWITCH.confirm primitive is used to inform the next higher layer of the initiating device whether the channel switch notification command is transmitted successfully.

The semantics of this primitive are:

MLME-CHANNELSWITCH.confirm 	(

status,

DeviceAddrMode,

DeviceAddress

)

The primitive parameters are defined in Table 44b.

Table 44b—MLME-CHANNELSWITCH.confirm parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| status | Enumeration | SUCCESS, TRANSACTION_OVERFLOW, TRANSACTION_EXPIRED, NO_ACK, CHANNEL_ACCESS_FAILURE, COUNTER_ERROR, FRAME_TOO_LONG, UNAVAILABLE_KEY, UNSUPPORTED_SECURITY, INVALID_PARAMETER | The status of the transmission of channel switch notification command. |
| DeviceAddrMode | Enumeration | SHORT_ADDRESS, EXTENDED_ADDRESS | The addressing mode of the device which has been instructed to change the operating channel at a specific time by its PAN coordinator. For MBANS devices, the default value is EXTENDED_ADDRESS |
| DeviceAddress | Device address | As specified by the DeviceAddrMode parameter | The address of the device which has been instructed to change the operating channel at a specific time. |

This primitive returns a status of either SUCCESS, indicating that the channel switch notification is successful, or the appropriate status parameter value indicating the reason for failure.

MLME-CHANNELSWITCH.indication

The MLME-CHANNELSWITCH.indication primitive is used to indicate the reception of a channel switch notification command.

The semantics of this primitive are:

MLME-CHANNELSWITCH.indication  	(

DeviceAddrMode,

DeviceAddress,

ChannelNumber,

ChannelPage,

NewPANID,

CoordinatorAddress,

RemainingTime,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44c.

Table 44c—MLME-CHANNELSWITCH.indication parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| DeviceAddressMode | Enumeration | SHORT_ADDRESS, EXTENDED_ADDRESS | The addressing mode of the device transmitting the channel switch notification command. For MBANS devices, the default value is EXTENDED_ADDRESS. |
| DeviceAddress | Device address | As specified by the DeviceAddrMode parameter | The address of the device transmitting the channel switch notification command. |
| ChannelNumber | Integer | Any valid channel number | As defined in 5.3.10.3. |
| ChannelPage | Integer | Any valid channel page | As defined in 5.3.10.4. |
| NewPANID | Integer | As defined in 5.3.10.2.1.1 | As defined in 5.3.10.2.1.1. |
| CoordinatorAddress | Device address | As defined in 5.3.10.2.1.2 | As defined in 5.3.10.2.1.2. |
| RemainingTime | Integer | As defined in 5.3.10.2.1.3 | As defined in 5.3.10.2.1.3. |
| SecurityLevel | Integer | As defined in Table 48 | As defined in Table 48. |
| KeyIdMode | Integer | As defined in Table 48 | As defined in Table 48. |
| KeySource | Set of octets | As defined in Table 48 | As defined in Table 48. |
| KeyIndex | Integer | As defined in Table 48 | As defined in Table 48. |

Insert:

Periodic GTS management primitives

These primitives are used to request and maintain periodic GTSs.

MLME-PERIODIC-GTS.request

The MLME-PERIODIC-GTS.request primitive allows a device to send a request to the PAN coordinator to allocate a new periodic GTS or to deallocate an existing periodic GTS. This primitive is also used by the PAN coordinator to initiate a periodic GTS deallocation.

The semantics of this primitive are:

MLME-PERIODIC-GTS.request 	(

PeriodicGTSCharacteristics,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44d.

Table 44d—MLME-PERIODIC-GTS.request parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| PeriodicGTSCharacteristics | Periodic GTS characteristics | As defined in 5.3.9.3 | The characteristics of the periodic GTS request. |
| SecurityLevel | Integer | As defined in Table 46 | As defined in Table 46. |
| KeyIdMode | Integer | As defined in Table 46 | As defined in Table 46. |
| KeySource | Set of octets | As defined in Table 46 | As defined in Table 46. |
| KeyIndex | Integer | As defined in Table 46 | As defined in Table 46. |

On receipt of the MLME-PERIODIC-GTS.request primitive by a device, the MLME of the device performs either the periodic GTS request procedure, as described in 5.1.7.2, or the periodic GTS deallocation procedure, as described in 5.1.7.4, depending on the value of the PeriodicGTSCharacteristics field.

MLME-PERIODIC-GTS.confirm

The MLME-PERIODIC-GTS.confirm primitive reports the results of a request to allocate a new periodic GTS or to deallocate an existing periodic GTS.

The semantics of this primitive are:

MLME-GTS.confirm	(

PeriodicGTSCharacteristics,

status

)

The primitive parameters are defined in Table 44e.

Table 44e—MLME-PERIODIC-GTS.confirm parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| PeriodicGTSCharacteristics | Periodic GTS characteristics | As defined in 5.3.9.3 | The characteristics of the periodic GTS request. |
| status | Enumeration | SUCCESS, DENIED, NO_SHORT_ADDRESS, CHANNEL_ACCESS_FAILURE, NO_ACK, NO_DATA, COUNTER_ERROR, FRAME_TOO_LONG, UNAVAILABLE_KEY, UNSUPPORTED_SECURITY, INVALID_PARAMETER. | The status of the periodic GTS request. |

If the request to allocate or deallocate a periodic GTS is successful, this primitive will return a status of SUCCESS and the Characteristics Type field of the PeriodicGTSCharacteristics parameter will have the value of one or zero, respectively. Otherwise, the status parameter will indicate the appropriate error code, as defined in 5.1.7.2 or 5.1.7.4.

If macShortAddress is equal to 0xfffe or 0xffff, the device is not permitted to request a periodic GTS and the status parameter will be set to NO_SHORT_ADDRESS.

MLME-PERIODIC-GTS.indication

The MLME-PERIODIC-GTS.indication primitive indicates that a periodic GTS has been allocated or that a previously allocated periodic GTS has been deallocated

The semantics of this primitive are:

MLME-PERIODIC-GTS.indication	(

DeviceAddress,

PeriodicGTSCharacteristics,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44f.

Table 44f—MLME-PERIODIC-GTS.indication parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| DeviceAddress | Device address | 0x0000–0xfffd | The short address of the device that has been allocated or deallocated a periodic GTS. |
| PeriodicGTSCharacteristics | Periodic GTS characteristics | As defined in 5.3.9.3 | The characteristics of the periodic GTS request. |
| SecurityLevel | Integer | As defined in Table 48 | If the primitive is generated when a periodic GTS deallocation is initiated by the PAN coordinator itself, the security level to be used is set to 0x00. If the primitive is generated whenever a periodic GTS is allocated or deallocated following the reception of a periodic GTS request command, then it is as defined in Table 48. |
| KeyIdMode | Integer | As defined in Table 48 | If the primitive is generated when a periodic GTS deallocation is initiated by the PAN coordinator itself, the security level to be used is set to 0x00. If the primitive is generated whenever a periodic GTS is allocated or deallocated following the reception of a periodic GTS request command, then it is as defined in Table 48. |
| KeySource | Set of octets | As defined in Table 48 | If the primitive is generated when a periodic GTS deallocation is initiated by the PAN coordinator itself, the security level to be used is set to 0x00. If the primitive is generated whenever a periodic GTS is allocated or deallocated following the reception of a periodic GTS request command, then it is as defined in Table 48. |
| KeyIndex | Integer | As defined in Table 48 | If the primitive is generated when a periodic GTS deallocation is initiated by the PAN coordinator itself, the security level to be used is set to 0x00. If the primitive is generated whenever a periodic GTS is allocated or deallocated following the reception of a periodic GTS request command, then it is as defined in Table 48. |

The value of the Characteristics Type field, as defined in 5.3.9.3, in the PeriodicGTSCharacteristics parameter indicates if the periodic GTS has been allocated or if a periodic GTS has been deallocated.

Grant association proxy notification primitives

These primitives are used when a device requests a grant of association proxy.

MLME-GRANTASSOCIATIONPROXY.request

The MLME-GRANTASSOCIATIONPROXY.request primitive is used by a FFD to request the grant of an association proxy with a coordinator.

The semantics of this primitive are:

MLME-GRANTASSOCIATIONPROXY.request	(

ChannelNumber,

ChannelPage,

CoordAddressMode,

CoordPANId,

CoordAddress,

NumberOfDevices,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44g.

Table 44g—MLME-GRANTASSOCIATIONPROXY.request parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| ChannelNumber | Integer | Any valid channel number | The channel number on which to attempt association proxy. |
| ChannelPage | Integer | Any valid channel page | The channel page on which to attempt association proxy. |
| CoordAddressMode | Enumeration | SHORT_ADDRESS, EXTENDED_ADDRESS | The coordinator addressing mode for this primitive and subsequent MPDU. |
| CoordPANId | Integer | 0x0000–0xffff | The identifier of the PAN with which to associate. |
| CoordAddress | Device address | As specified by the CoordAddrMode parameter | The address of the coordinator with which to associate. |
| NumberOfDevices | Integer | 0xa0-0xbf | Specifies the number of devices for which a grant association proxy is requested. The number is offset by 0xa0. |
| SecurityLevel | Integer | As defined in Table 46 | As defined in Table 46. |
| KeyIdMode | Integer | As defined in Table 46 | As defined in Table 46. |
| KeySource | Set of octets | As defined in Table 46 | As defined in Table 46. |
| KeyIndex | Integer | As defined in Table 46 | As defined in Table 46. |

MLME-GRANTASSOCIATIONPROXY.indication

The MLME-GRANTASSOCIATIONPROXY.indication primitive is used to indicate the reception of a grant association proxy request command.

The semantics of this primitive are:

MLME-GRANTASSOCIATIONPROXY.indication	(

DeviceAddress,

NumberOfDevices,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44i

Table 44i—MLME-GRANTASSOCIATIONPROXY.indication parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| DeviceAddress | Device address | An extended IEEE address | The address of the device requesting a grant of association proxy. |
| NumberOfDevices | Integer | 0xa0-0xbf | Specifies the number of devices for which an association proxy is requested. The number of devices is offset by 0xa0. |
| SecurityLevel | Integer | As defined in Table 48 | As defined in Table 48. |
| KeyIdMode | Integer | As defined in Table 48 | As defined in Table 48. |
| KeySource | Set of octets | As defined in Table 48 | As defined in Table 48. |
| KeyIndex | Integer | As defined in Table 48 | As defined in Table 48. |

When the next higher layer of a coordinator receives the MLME-GRANTASSOCIATIONPROXY.indication primitive, the coordinator determines whether to accept or reject the unassociated devices using an algorithm outside the scope of this standard.

MLME-GRANTASSOCIATIONPROXY.response

The MLME-GRANTASSOCIATIONPROXY.response primitive reports the results of an MLME-GRANTASSOCIATIONPROXY.request primitive.

The semantics of this primitive are:

MLME-GRANTASSOCIATIONPROXY.response 	(

DeviceAddress,

NumberAllocatedShortAddresses,

AssocShortAddress,

status,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44j.

Table 44j—MLME-GRANTASSOCIATIONPROXY.response parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| DeviceAddress | Device address | An extended 64 bit IEEE address | The address of the device requesting association. |
| NumberAllocatedShortAddresses | Integer | 0xa0-0xbf | The number of allocated short addresses. |
| AssocShortAddress | Integer | 0x0000–0xfffd | The short device address allocated by the coordinator on successful grant association proxy. This parameter is repeated for all of the allocated short addresses. |
| status | Enumeration | As defined in 5.3.2.3 | The status of the grant association proxy attempt. |
| SecurityLevel | Integer | As defined in Table 46 | As defined in Table 46. |
| KeyIdMode | Integer | As defined in Table 46 | As defined in Table 46. |
| KeySource | Set of octets | As defined in Table 46 | As defined in Table 46. |
| KeyIndex | Integer | As defined in Table 46 | As defined in Table 46. |

When the MLME of a coordinator receives the MLME-GRANTASSOCIATIONPROXY.response primitive, it generates a grant association proxy response command, as described in 5.3.12 and attempts to send it to the device requesting grant association proxy, as described in 5.1.3.3.

MLME-GRANTASSOCIATIONPROXY.confirm

The MLME-GRANTASSOCIATIONPROXY.confirm primitive is used to inform the next higher layer of the initiating device whether its request to grant an association proxy is successful or unsuccessful.

The semantics of this primitive are:

MLME-GRANTASSOCIATIONPROXY.confirm	(

NumberAllocatedShortAddresses,

AssocShortAddress,

status,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44k.

Table 44k—MLME-GRANTASSOCIATIONPROXY.confirm parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| NumberAllocatedShortAddresses | Integer | 0xa0-0xbf | The number of allocated short addresses. |
| AssocShortAddress | Integer | 0x0000-0xfffd | The short device address allocated by the coordinator on successful grant association proxy. This parameter is repeated for all of the allocated device short addresses. |
| status | Enumeration | The value of the status field of the association response command, as defined in 5.3.2.3, SUCCESS, CHANNEL_ACCESS_FAILURE, NO_ACK, NO_DATA, COUNTER_ERROR, FRAME_TOO_LONG, IMPROPER_KEY_TYPE, IMPROPER_SECURITY_LEVEL, SECURITY_ERROR, UNAVAILABLE_KEY, UNSUPPORTED_LEGACY, UNSUPPORTED_SECURITY, INVALID_PARAMETER | The status of the association attempt. |
| SecurityLevel | Integer | As defined in Table 48 | As defined in Table 48. |
| KeyIdMode | Integer | As defined in Table 48 | As defined in Table 48. |
| KeySource | Set of octets | As defined in Table 48 | As defined in Table 48. |
| KeyIndex | Integer | As defined in Table 48 | As defined in Table 48. |

If the grant association proxy request is successful, then the status parameter shall be set to SUCCESS. Otherwise, the status parameter shall be set to indicate the type of failure.

Grant association proxy notification primitives

These primitives are used when a device requests a grant of association proxy.

MLME-ASSOCIATIONPROXY.request

The MLME-ASSOCIATIONPROXY.request primitive is used by a FFD to update the coordinator with the extended address and Capability Information of a device that has been associated using the grant associated proxy command 5.3.11.

The semantics of this primitive are:

MLME-ASSOCIATIONPROXY.request	(

CoordAddressMode,

CoordPANId,

CoordAddress,

AssocShortAddress,

DeviceAddress,

CapabilityInformation,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44l.

Table 44l—MLME-ASSOCIATIONPROXY.request parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| CoordAddressMode | Enumeration | SHORT_ADDRESS, EXTENDED_ADDRESS | The addressing mode of the coordinator to which the association proxy is intended. |
| CoordPANId | Integer | 0x0000–0xffff | The PAN identifier of the coordinator to which the association proxy is intended. |
| CoordAddress | Device address | As specified by the CoordAddrMode parameter | The address of the coordinator to which the association proxy in intended. |
| AssocShortAddress | Integer | 0x0000–0xfffd | The short device address allocated by the coordinator on successful association. |
| DeviceAddress | Device address | An extended address | The address of the device requesting an association proxy. |
| CapabilityInformation | Bitmap | As defined in 5.3.1.2 | The operational capabilities of the device requesting an association proxy. |
| SecurityLevel | Integer | As defined in Table 46 | As defined in Table 46. |
| KeyIdMode | Integer | As defined in Table 46 | As defined in Table 46. |
| KeySource | Set of octets | As defined in Table 46 | As defined in Table 46. |
| KeyIndex | Integer | As defined in Table 46 | As defined in Table 46. |

The AssocShortAddress is the address that is allocated to a device during the grant association proxy request. The DeviceAddress and Capability Information are for this device.

MLME-ASSOCIATIONPROXY.indication

The MLME-ASSOCIATIONPROXY.indication primitive is used to inform the next higher layer of the coordinator of the extended address and Capability Information for a device that is allocated a short address during the grant association proxy request.

The semantics of this primitive are:

MLME-ASSOCIATIONPROXY.indication	(

CoordAddressMode,

CoordPANId,

CoordAddress,

AssocShortAddress,

DeviceAddress,

CapabilityInformation,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44m.

Table 44m—MLME-ASSOCIATIONPROXY.indication parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| CoordAddressMode | Enumeration | SHORT_ADDRESS, EXTENDED_ADDRESS | The addressing mode of the coordinator to which the association proxy is intended. |
| CoordPANId | Integer | 0x0000–0xffff | The PAN identifier of the coordinator to which the association proxy is intended. |
| CoordAddress | Device address | As specified by the CoordAddrMode parameter | The address of the coordinator to which the association proxy in intended. |
| AssocShortAddress | Integer | 0x0000–0xfffd | The short device address allocated by the coordinator on successful association. |
| DeviceAddress | Device address | An extended address | The address of the device requesting an association proxy. |
| CapabilityInformation | Bitmap | As defined in 5.3.1.2 | The operational capabilities of the device requesting an association proxy. |
| SecurityLevel | Integer | As defined in Table 48 | As defined in Table 48. |
| KeyIdMode | Integer | As defined in Table 48 | As defined in Table 48. |
| KeySource | Set of octets | As defined in Table 48 | As defined in Table 48. |
| KeyIndex | Integer | As defined in Table 48 | As defined in Table 48. |

MLME-ASSOCIATIONPROXY.confirm

The MLME-ASSOCIATIONPROXY.confirm primitive reports the results of an association proxy command to update the coordinator with the associated device-specific information.

The semantics of this primitive are:

MLME-ASSOCIATIONPROXY.confirm	(

AssocShortAddress,

DeviceAddress,

status

)

The primitive parameters are defined in Table 44n.

Table 44n—MLME.ASSOCIATIONPROXY.confirm parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| AssocShortAddress | Integer | 0x0000–0xfffd | The short device address allocated by the coordinator on successful association. |
| DeviceAddress | Device address | An extended address | The address of the device requesting an association proxy. |
| status | Enumeration | The value of the status field of the association response command, as defined in 5.3.2.3, SUCCESS, DENIED, NO_SHORT_ADDRESS, CHANNEL_ACCESS_FAILURE, NO_ACK, NO_DATA, COUNTER_ERROR, FRAME_TOO_LONG, UNAVAILABLE_KEY, UNSUPPORTED_SECURITY, INVALID_PARAMETER | The status of the association proxy request. |

If the association proxy request is successful, then the status parameter shall be set to SUCCESS. Otherwise, the status parameter shall be set to indicate the type of failure.

Coordinator switch notification primitives

These primitives are used when a device requests a coordinator switch.

MLME-COORDINATOR-SWITCH.request

The MLME-COORDINATOR-SWITCH.request primitive is used by a PAN coordinator to request a coordinator switch from other coordinators.

The semantics of this primitive are:

MLME-COORDINATOR-SWITCH.request	(

ChannelNumber,

ChannelPage,

SrcAddrMode,

DstAddrMode,

NumberOfDevices,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44o.

Figure 44o–MLME-COORDINATOR-SWITCH.request parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| ChannelNumber | Integer | Any valid channel number | The channel number on which to attempt a coordinator switch. |
| ChannelPage | Integer | Any valid channel page | The channel page on which to attempt a coordinator switch. |
| SrcAddrMode | Enumeration | SHORT_ADDRESS, EXTENDED_ADDRESS | The source addressing mode for this MPDU. |
| DstAddrMode | Enumeration | SHORT_ADDRESS, EXTENDED_ADDRESS | The destination addressing mode for this MPDU. |
| NumberOfDevices | Integer | 0x00-0xff | Specifies the number of devices that the coordinator wishes to switch to another coordinator. |
| SecurityLevel | Integer | As defined in Table 46 | As defined in Table 46. |
| KeyIdMode | Integer | As defined in Table 46 | As defined in Table 46. |
| KeySource | Set of octets | As defined in Table 46 | As defined in Table 46. |
| KeyIndex | Integer | As defined in Table 46 | As defined in Table 46. |

On receipt of the MLME-COORDINATOR-SWITCH.request primitive, the MLME of a coordinator shall first update the appropriate PHY and MAC PIB attributes, and then generates a coordinator switch command as defined in 5.3.15.

MLME-COORDINATOR-SWITCH.indication

The MLME-COORDINATOR-SWITCH.indication primitive is used to indicate the reception of a coordinator switch request command.

The semantics of this primitive are as follows:

MLME-COORDINATOR-SWITCH.indication	(

CoordPANId,

DeviceAddress,

NumberOfDevices,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44p

Table 44p–MLME-COORDINATOR-SWITCH.indication parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| CoordPANId | Integer | 0x0000–0xfffe | The PAN identifier of the coordinator requesting a coordinator switch. |
| DeviceAddress | Device address | An extended address | The address of the coordinator requesting a coordinator switch |
| NumberOfDevices | Integer | 0x00-0xff | Specifies the number of devices for which a coordinator switch is being requested. |
| SecurityLevel | Integer | As defined in Table 48 | As defined in Table 48. |
| KeyIdMode | Integer | As defined in Table 48 | As defined in Table 48. |
| KeySource | Set of octets | As defined in Table 48 | As defined in Table 48. |
| KeyIndex | Integer | As defined in Table 48 | As defined in Table 48. |

When the next higher layer of a coordinator receives the MLME-COORDINATOR-SWITCH.indication primitive, the next higher layer shall determine whether it has the resources to be able to accept the association of additional devices using an algorithm outside the scope of this standard.

MLME-COORDINATOR-SWITCH.response

The MLME-COORDINATOR-SWITCH.response primitive reports the results of an MLME-COORDINATOR-SWITCH.request primitive.

The semantics of this primitive are as follows:

MLME-COORDINATOR-SWITCH.response 	(

CoordPANId,

DeviceAddress,

NumberOfDevices,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44q.

Table 44q–MLME-COORDINATOR-SWITCH.response parameters

| Name | type | Valid range | Description |
| --- | --- | --- | --- |
| CoordPANId | Integer | 0x0000–0xfffe | The PAN identifier of the coordinator requesting a coordinator switch. |
| DeviceAddress | Device address | An extended address | The address of the coordinator requesting a coordinator switch. |
| NumberOfDevices | Integer | 0x00–0xff | Confirms the number of devices for which a coordinator switch is being requested. |
| SecurityLevel | Integer | As defined in Table 46 | As defined in Table 46. |
| KeyIdMode | Integer | As defined in Table 46 | As defined in Table 46. |
| KeySource | Set of octets | As defined in Table 46 | As defined in Table 46. |
| KeyIndex | Integer | As defined in Table 46 | As defined in Table 46. |

When the MLME of a coordinator receives the MLME-COORDINATOR-SWITCH.response primitive, it shall generate a coordinator response command as described in 5.3.16 in response to a broadcast coordinator switch request and in response to a unicast coordinator switch request command as described in 5.1.2.6.

MLME-COORDINATOR-SWITCH.confirm

The MLME-COORDINATOR-SWITCH.confirm primitive is used to inform the next higher layer of the initiating device whether its request for coordinator switch is successful or unsuccessful.

The semantics of this primitive are:

MLME-COORDINATOR-SWITCH.confirm	(

CoordPANId,

DeviceAddress,

NumberOfDevices,

status,

SecurityLevel,

KeyIdMode,

KeySource,

KeyIndex

)

The primitive parameters are defined in Table 44r.

Table 44r–MLME-COORDINATOR-SWITCH.confirm parameters

| Name | Type | Valid range | Description |
| --- | --- | --- | --- |
| CoordPANId | Integer | 0x0000–0xfffe | The PAN identifier of the new coordinator. |
| DeviceAddress | Device address | An extended address | The address of the new coordinator. |
| NumberOfDevices | Integer | 0x00–0xff | Confirms the number of devices for which a coordinator switch is being requested. If this value is 0x00 then the coordinator switch request was unsuccessful. |
| status | Enumeration | SUCCESS, CHANNEL_ACCESS_FAILURE, NO_ACK, NO_DATA, COUNTER_ERROR, FRAME_TOO_LONG, IMPROPER_KEY_TYPE, IMPROPER_SECURITY_LEVEL, SECURITY_ERROR, UNAVAILABLE_KEY, UNSUPPORTED_LEGACY, UNSUPPORTED_SECURITY, INVALID_PARAMETER | The status of the association attempt. |
| SecurityLevel | Integer | As defined in Table 48 | As defined in Table 48. |
| KeyIdMode | Integer | As defined in Table 48 | As defined in Table 48. |
| KeySource | Set of octets | As defined in Table 48 | As defined in Table 48. |
| KeyIndex | Integer | As defined in Table 48 | As defined in Table 48. |

If the coordinator switch request is successful, then the status parameter will be set to SUCCESS. Otherwise, the status parameter will be set to indicate the type of failure.

MAC data service

MAC constants and PIB attributes

MAC constants

MAC PIB attributes

Insert as shown in Table 52:

Table 52–MAC PIB attributes

| Attribute | Type | Range | Description | Default |
| --- | --- | --- | --- | --- |
| macPeriodicGTSPermit* | Boolean | TRUE, FALSE | TRUE if the PAN coordinator is to accept Periodic GTS requests. FALSE otherwise. | TRUE |

Security

General PHY requirements

General requirements and definitions

Iinsert in 8.1:

O-QPSK PHY: direct sequence spread spectrum (DSSS) PHY employing offset quadrature phase¬shift keying (O-QPSK) modulation, operating in the 780 MHz bands, 868 MHz, 915 MHz, 2380 MHz, and 2450 MHz, as defined in clause 10.

Operating frequency range

Insert as shown in table 66:

Table 66–Frequency bands and data rates

| PHY (MHz) | Frequency band (MHz) | Spreading parameters | Data parameters |
| --- | --- | --- | --- |
|  |  | Chip rate (kchip/s) | Modulation | Bit rate (kb/s) | Symbol rate (ksymbol/s) | Symbols |
| 2380 DSSS | 2360-2400 | 2000 | O-QPSK | 250 | 62.5 | 16-ary Orthogonal |

Channel assignments

Channel numbering for 780 MHz band

Channel numbering for 868 MHz, 915 MHz, and 2450 MHz bands

Channel numbering for 950 MHz band

Channel numbering for CSS PHY

Channel numbering for UWB PHY

Insert after 8.1.2.5:

Channel numbering for 2380 MHz band

For channel page 7, 15 channels numbered zero to fourteen are available across the 2380 MHz band. The center frequencies of these channels are defined as follows:

Fc = 2363 + 5 k in MHz, for k = 0, 1, …, 6

Fc = 2367 + 5 (k – 7) in MHz, for k = 7, 8, …, 13

and Fc = 2395 in MHz, for k = 14

where

k is the channel number.

PHY services

O-QPSK PHY

PPDU format

Modulation and spreading

Data rate

Change and insert in the first paragraph of 10.2.1:

The data rate of the O-QPSK PHY shall be 250 kb/s when operating in the 2450 MHz, 915 MHz, or 780 MHz, or 2380 MHz bands and shall be 100 kb/s when operating in the 868 MHz band.

Reference modulator diagram

Bit-to-symbol mapping

Symbol-to-chip mapping

Insert in the first paragraph of 10.2.4:

In the 2450 MHz and 2380 MHz bands, each data symbol shall be mapped into a 32-chip PN sequence as specified in Table 73. The PN sequences are related to each other through cyclic shifts and/or conjugation (i.e., inversion of odd-indexed chip values).

Insert in the title of Table 73:

Table 73—Symbol-to-chip mapping for the 2450 MHz and 2380 MHz bands

O-QPSK modulation

Insert in the first paragraph of 10.2.5:

The chip sequences representing each data symbol are modulated onto the carrier using O-QPSK with half-sine pulse shaping. Even-indexed chips are modulated onto the in-phase (I) carrier, and odd-indexed chips are modulated onto the quadrature-phase (Q) carrier. In the 2450 MHz and 2380 MHz bands, each data symbol is represented by a 32-chip sequence, and so the chip rate is 32 times the symbol rate. In the 915 MHz, 868 MHz, and 780 MHz bands, each data symbol is represented by a 16-chip sequence, and so the chip rate is 16 times the symbol rate. To form the offset between I-phase and Q-phase chip modulation, the Q-phase chips shall be delayed by Tc with respect to the I-phase chips, as illustrated in Figure 70, where Tc is the inverse of the chip rate.

Pulse shape

Change and insert in the first paragraph of 10.2.6:

In the 2450 MHz, 915 MHz, and 868 MHz, and 2380 MHz bands, the half-sine pulse shape is used to represent each baseband chip and is given by:

Chip transmission order

Insert in 10.2.7:

During each symbol period, the least significant chip, c0, is transmitted first and the most significant chip, either c31, for the 2450 MHz and 2380 MHz bands, or c15, for the 915 MHz, 868 MHz, and 780 MHz bands, is transmitted last.

O-QPSK PHY RF requirements

Operating frequency range

Insert as shown in 10.3.1:

— 779–787 MHz

— 868.0–868.6 MHz

— 902–928 MHz

— 2360–2400 MHz

— 2400.0–2483.5 MHz

Transmit power spectral density (PSD) mask

Insert in the third paragraph of 10.3.2:

When operating in the 2380 MHz and 2450 MHz bands, the transmitted spectral products shall be less than the limits specified in Table 76. For both relative and absolute limits, average spectral power shall be measured using a 100 kHz resolution bandwidth. For the relative limit, the reference level shall be the highest average spectral power measured within ±1 MHz of the carrier frequency.

Insert in the title of table 76:

Table 76—2380 MHz and 2450 MHz bands O-QPSK transmit PSD limits

Symbol rate

Insert as shown in 10.3.3:

The O-QPSK PHY symbol rate shall be 25 ksymbol/s when operating in the 868 MHz band and 62.5 ksymbol/s when operating in the 780 MHz, 915 MHz, 2380 MHz, or 2450 MHz bands with an accuracy of ±40 ppm.

Receiver sensitivity

Receiver interference rejection

Insert in the first paragraph of 10.3.5:

This subclause applies only to the 780 MHz, 915 MHz, 2380 MHz, and 2450 MHz bands as there is only one channel available in the 868 MHz band.

Insert in the title of table 77:

Table 77—Minimum receiver interference rejection requirements for the 780 MHz, 915 MHz, 2380 MHz, and 2450 MHz bands

Insert the following Annex I:

# Annex I

(informative)

Channel bitmap

I.1 General

To assist with coexistence to the primary users of the 2360-2400 MHz band, some parts of this band may be excluded from use by MBANS devices. Those parts of the band that are excluded from use may change from time to time. Regulation requires that the operation of MBANS hub devices have access to a mechanism that provides information on which parts of the band cannot be used. The description of this mechanism is beyond the scope of this annex. To assist other MBANS devices such as sensors that do not have direct access to this mechanism, a bitmap may be used by a hub device to indicate those channels that are allowed to be used by MBANS devices. The bitmap only applies to those channels that are in the 2360-2390 MHz band (channels 0-5 and 7-12). The channels in the 2390-2400 MHz band (channels 6, 13 and 14) are always available for use.

I.2 Channel bitmap

A bitmap of those channels that are available for use may be inserted into the beacon payload. The bitmap shall be formatted as shown in Table I.1.

Table I.1–Channel bitmap

| Bits: 0-11 | Bit 12-22 | Bit 23 |
| --- | --- | --- |
| Channel bitmap | Valid time | Reserved |

The bit-to-channel numbering of the channel bitmap field is shown in Table I.2. The channel numbering is described in 8.1.2.6.

Table I.2–Bit-to-channel numbering

| Bit | Channel, k |
| --- | --- |
| 0 | 0 |
| 1 | 1 |
| 2 | 2 |
| 3 | 3 |
| 4 | 4 |
| 5 | 5 |
| 6 | 7 |
| 7 | 8 |
| 8 | 9 |
| 9 | 10 |
| 10 | 11 |

The bits of the channel bitmap field shall be set to one when the corresponding channels are available for use and the bits of the channel bitmap field shall be set to zero when the corresponding channels are not available for use.

The Valid time field indicates the duration in minutes of the availability of the channel bitmap.

# Footnotes

IEEE publications are available from the Institute of Electrical and Electronics Engineers, 445 Hoes Lane, P.O. Box 1331, Piscataway, NJ 08855-1331, USA (http://standards.ieee.org).

MBANS hub devices are 802.15.4 PAN coordinators.